
**PeakForce SECM with Bio-Logic SP-300 Potentiostat**

**Weilai Yu (Caltech)**

2018.12.29


**Preliminaries:**

1. For use of a Bio\_logic bipotentiostat in place of a CHI760E, follow the Bruker PeakForce SECM manual to set up all settings except skipping the CHI760E bipotentiostat in that manual.
2. To guarantee the noise level of ~10 pA at 10 nA current range with 1 kHz filter, both the ultra-low current probes (or amplifiers) of the two low-current cables from SP-300 Potentiostas **MUST** be inside the AFM hood, as shown in Figure 1. The AFM hood itself is a effective Faraday cage.



**Figure 1.** Low Current Probes (or Amplifiers) located inside the AFM hood, an effective Faraday cage. Cables are routed through a sealed clamp slot.


1. It is not clear what the SP-300 bipotentiostat bandwidth is when the current is 1–10 nA. A the filter setting of 5 Hz, is too slow for AFM imaging. From our testing, either 1 kHz or 50 kHz has similar noise level. This may suggest that the bandwidth might be below 1 kHz.
2. The noise level was tested using CA setting with setting time interval at 0.03 s, the filter at 1kHz, and the current range at 10 nA. The noise level is about 10 pA or 20 pA peak-to-peak. The standard deviation of the noise is ~9 pA. This is shown in Figure 2.



**Figure 2.** Typical current noise level (~10 pA) in CA mode. The noise level was tested using CA setting with setting time interval at 0.03 s, the filter at 1kHz, and the current range at 10 nA.

**Installation:**


1. Connect the two low-current cables to the SP-300 Potentiostats and insert the two cables into the hood of the Dimension Icon through the holes at its bottom; Channel 1 and 2 cables are shown on the left and right; respectively. The potentiostat cables can be fixed to appropriate position using tape; Low-current probes (blue ports) should always stay inside the AFM Hood during measurement to reduce the noise level of the gained current signal;



**Figure 3.** Hardware configuration between the Bio-Logic SP300 potentiostat and the Dimension Icon AFM

1. For the use of the Bio-Logic as a bipotentiostat you **must** configure the instrument


to use the counter electrode (CE) as ground (CE to Ground). To do this



**Figure 4:** CE to ground connection mode for using SP-300 as Bipotentiostat **.**



* 1. In the EClab Express software under Hardware configuration choose CE to Ground or in “Advanced Settings”, select the “CE-to-Ground” option.
  2. **Disconnect the cables from the cell.**,
  3. Connect the P1 and S1 (red) leads to the working electrode, Figure 4.
  4. Connect the S2 (black) lead to the reference electrode
  5. Connect the GROUND (black) and S3 (blue) leads to the counter electrode
  6. The P2 dark blue cable is not be used (open-circuit);
  7. Using only the alligator clips, these cables can be connected to the multi-channel BNC converters provided by Bruker and finally to the small electrochemical spring terminal on the SECM chuck;

1. The female DB9 connector can be used to interconnect the SP300 potentiostat with the AFM controller by use of the DB9 to multi-BNC converter. Two of the outputs from the converter are labeled E and I. Note that the E and I monitors have the opposite polarity to the real measured values i.e., if the analog output say -2V, it’s actually +2V at the cell.
2. For SECM experiments, connect the Bio-Logic channel 1 S1 and P1 to the working electrode to the nano-tip working electrode and channel 2 (S1 and P1) to the sample. To allow the AFM controller to monitor the currents and potentials for the tip and sample you need to connect the outputs E and I of the Bio-Logic to the Icon controller. Connect Channel 1 outputs E and I from the Bio-Logic to the AFM controller using the “Low Frequency Analog” panel inputs 1 and 2. To moitor the samplee current (or working electrode 2 in a 4-electrode configuration), connect Channel 2 output I to and input to Lock-In 3 input.




1. **Figure 5.** The BNC connection from the SP300 potentiostat (E/I signals) to the AFM controller (Input 1/2) so that the electrochemical signals can be directly recorded on the NanoScope Software. To monitor the sample current (or working electrode 2 in a 4-electrode configuration), connect channel 2 to Lock-In 3 input.
2. Before starting EC experiments, in the EC-Lab “Advanced Settings”, make sure to select the “CE-to-Ground” option chosen for a bi-potentiostat. The band width can be set at 1kHz to allow consistent data collections with scanning speed of the tip; In selecting experiment, one can either choose to select to use a normal single-channel experiment (CA/CV), or a combined bi-potetiostat experiments (CV-CA, CA-CA, CP-CA) by choosing the Channels 1/2. Before starting an experiment, a dialog box may appear asking for the change of the electrode configurations, be sure to select “No” and keep the configuration at “CE to ground”.


**Figure 6**. The connection of Bio-Logic Cables and Clips and the four-terminal converter inside the hood of the Dimension Icon




**7**

**Figure 8.** When selecting either a single-channel or a bi-channel experiment, be sure to select “CE to ground” as the electrode connection





1. During the scan, the noise level in CA mode with a time interval what is the time interval the time per point, the lines in the plot or between the dotted lines? of 0.03 s and band width of 1kHz was found to be ~10 pA (Figure 9).

**9**




1. When recording electrochemical data in the NanoScope Software using the XY data capture, the current range selected is the same EC-Lab and NanoScope Software. For example, if the I-range is “1 nA” in the EC-Lab, “1 nA/V” should be selected in the NanoScope Software.
2. Figure 9 shows a typical CV graph of the tip working electrode in the NanoScope Software recorded by the AFM controller.






**Figure 10.** Typical CV curve of the tip electrode recorded in the Bio-Logic software and in the strip chart of the NanoScope software. Note that the polarity is reversed which can be adjusted in the NanoScope software.

1. [Warning] During experiment, an obvious discharging behavior from the tip electrode at the beginning of each experiment was observed. This may be due to the ultra-small size the tip. This discharging behavior may impact the quality of the electrochemical data gained at the start of experiment. For example, The Figure below shows several continuous CV curves of the tip electrode. Due to the discharging behavior, the first CV was impacted and the CVs obtained later on appeared more well-defined and reasonable.



**Figure 10.** Typical charging behavior of the tip electrode at the beginning of experiment and its impact on the gained first CV curve

1. From here, following the PeakForce SECM manual to continue your experiment. Below is a screen show of SECM tip current profiles at different tip-sample distance. The line plots on the right from bottom are at nominal 30, 100, 200, 400, and 750 nm tip-sample distance. The noise level is about 10pA (std = 9.3 pA).



**Figure 11.** A typical SECM results using the Bio-Logic SP-300 bi-potentiostat.