SDK-S

User Manual

Important Changes & Compatibility	5

Introduction	5

Version	5

Installation	6

USB 3.0/2.0/1.1 Interface Spectrometers	11

USB Basic Functions	11

InitDevices	11

bwtekSetupChannel	12

bwtekReadEEPROMUSB	13

bwtekTestUSB	14

bwtekSetTimeUSB	15

bwtekSetTimeBase0USB	17

bwtekSetTimingsUSB	18

bwtekDataReadUSB	19

bwtekReadResultUSB	21

bwtekDataReadUSB1	24

bwtekGetTimeUSB	26

GetDeviceCount	27

GetUSBType	27

bwtekGetCCode (Obsolete)	28

GetCCode	28

bwtekCloseUSB	29

CloseDevices	29

Supplementary Functions	30

bwtekSmoothingUSB	30

bwtekConvertDerivativeDouble	31

bwtekPolyFit	32

bwtekPolyCalc	33

bwtekDataExport	34

bwtekSaveEEPROMChannel (Obsolete)	36

bwtekGetXaxisInverseByte	36

Temperature Readout – Limited Use	37

bwtekReadTemperature	37

USB 3.0 Interface Spectrometers	38

bwtekDSPDataReadUSB	38

bwtekFrameDataReadUSB	40

bwtekWriteBlockUSB	41

bwtekReadBlockUSB	42

bwtekEraseBlockUSB	43

AUX Port Functions 1 - Multi-Purpose TTL Output Functions	44

bwtekGetExtStatus	44

bwtekSetExtLaser	45

bwtekSetExtShutter	46

bwtekSetExtSync	47

bwtekGatedMode	48

bwtekSetExtPulse	50

Internal Trigger Mode	50

External Trigger Mode	52

AUX Port Functions 2---Shutter Functions:	55

bwtekShutterOpen (OBSOLETE)	55

bwtekShutterClose (OBSOLETE)	55

bwtekShutterControl	55

AUX Port Functions 3	57

bwtekGetExtStatus	57

bwtekGetTTLIn	58

bwtekSetTTLOut	59

bwtekGetAnalogIn	60

BTC261E & BTC262E Functions:	61

bwtekSetABGain (Obsolete)	61

bwtekSetABOffset (Obsolete)	63

bwtekGetABGain (Obsolete)	65

bwtekGetABOffset (Obsolete)	66

bwtekSetInGaAsMode	67

bwtekGetInGaAsMode	68

bwtekQueryTemperature	69

bwtekAccessDeltaTemp	70

bwtekAccessDeltaTemp1	71

bwtekWriteValue	73

bwtekReadValue	74

BTC262A &BTC263E Functions:	75

bwtekSetTimeUnitUSB	75

bwtekGetTimeUnitUSB	76

bwtekSetInGaAsMode	77

bwtekGetInGaAsMode	78

BTC261P & BTC262P & BTC264P Functions:	79

bwtekSetTimeUnitUSB (OBSOLETE)	79

bwtekGetTimeUnitUSB (OBSOLETE)	79

bwtekSetInGaAsMode	80

bwtekGetInGaAsMode	81

Appendix A: Model Descriptions	82

Appendix B: USB Interface Capabilities	83

Appendix C: Spectrometer Driver Files	84

Appendix D: Spectrometer’s USB Drivers File Destinations	85

Appendix E: SDK Function Groups	86

Appendix F: Default Parameters for bwtekTESTUSB	90

Appendix G: External IO Control	92

Appendix H: X-axis Reverse Settings	93

Appendix I: Dynamically Assigned Channel Numbers	95

# Important Changes & Compatibility

USB Device Channel Numbers for spectrometers can no longer be assigned manually.

Channel Numbers are now dynamically assigned, see function bwtekSetupChannel for further details.

For quick reference on the order in which to call the new functions in this SDK as they relate to the important changes for Dynamically Assigned USB Device Channel Numbers, reference Appendix: Dynamically Assigned Channel Numbers.

Compatibility: If you are using an older version of the SDK-S or SDK-L which allows for USB Device Channel Numbers to be manually assigned, it will Not be compatible with this newer version which uses Dynamically Assigned USB Device Channel Numbers.

# Introduction

B&W Tek, Inc. Software Developer’s Kit (SDK) is designed for customers who need to develop a custom program for their USB based BTC and/or BRC spectrometers for specific applications for 32-Bit and 64-Bit windows based software.

For RS232 Devices, contact B&W Tek, Inc. with your model spectrometer and serial number and reference RS232 Command Set, (Doc# 4000000008).

The SDK program’s .DLL files are compatible with any Windows XP, Windows 7 and Windows 8 32-Bit and 64-Bit Windows Operating Systems.

# Version

This document is based on .DLL version #:

- bwtekusb.dll (4.8.0.14)

# Installation

To use SDK-S you must first place specific files into folder on your computer.  Find the Operating System of your Windows-based PC and then find the correct Bit OS you have.  Follow the instructions for where to place the files on your PC.

- Locate the Operating System and Bit (32-Bit 64-Bit) for your computer.  Follow the instructions on where to place the files located into the SDK-S Package on your local computer.

- File and Folder Destination Folders:

  - Windows XP

    - Windows XP - 32-Bit

      - Windows XP – 32-Bit - USB 2.0 Capable Units:

- Copy “\32bit\Drivers\USB2\winxp\bwtekusb2.inf” ( c:\windows\inf\

- Copy “\32bit\Drivers\USB2\winxp\ fx2lp.cat” ( c:\windows\inf\

- Copy “\32bit\Drivers\USB2\winxp\*.spt” ( c:\windows\system32\drivers\bwtek\

- Copy “\32bit\Drivers\USB2\winxp\CyUSB.sys” ( c:\windows\system32\Drivers\

- Copy “\32bit\Drivers\CyUSB.dll” files ( c:\windows\system32\

- Copy “\32bit\DLL_32bit\bwtekusb.dll” ( the folder of your executable file.

      - Windows XP – 32-Bit - USB 3.0 Capable Units:

- Copy “\32bit\Drivers\USB3\winxp\bwtekusb3.inf” ( c:\windows\inf\

- Copy “\32bit\Drivers\ USB3\winxp\cyusb3.cat” ( c:\windows\inf\

- Copy “\32bit\Drivers\ USB3\winxp\*.spt” ( c:\windows\system32\drivers\bwtek\

- Copy “\32bit\Drivers\USB3\winxp\CyUSB3.sys” ( c:\windows\system32\Drivers\

- Copy “\32bit\Drivers\USB3\winxp\WdfCoInstaller01009.dl” ( c:\windows\system32\

- Copy “\32bit\Drivers\USB3\winxp\WdfCoInstaller01009.dll”( c:\windows\system32\Drivers\

- Copy “\32bit\Drivers\CyUSB.dll” files ( c:\windows\system32\

- Copy “\32bit\DLL_32bit\bwtekusb.dll” ( the folder of your executable file.

  - Windows 7

    - Windows 7 - 32-Bit

      - Windows 7 – 32-Bit - USB 2.0 Capable Units:

- Copy “\32bit\Drivers\USB2\win7_vista\bwtekusb2.inf” ( c:\windows\inf\

- Copy “\32bit\Drivers\USB2\win7_vista\fx2lp.cat” ( c:\windows\inf\

- Copy “\32bit\Drivers\USB2\win7_vista\*.spt” ( c:\windows\system32\drivers\bwtek\

- Copy “\32bit\Drivers\USB2\win7_vista\CyUSB.sys” ( c:\windows\system32\Drivers\

- Copy “\32bit\Drivers\CyUSB.dll” ( c:\windows\system32\

- Copy “\32bit\DLL_32bit\bwtekusb.dll” ( c:\windows\system32\

- Copy “\32bit\DLL_32bit\bwtekusb.dll” ( the folder of your executable file.

      - Windows 7 – 32-Bit - USB 3.0 Capable Units:

- Copy “\32bit\Drivers\USB3\win7\bwtekusb3.inf” ( c:\windows\inf\

- Copy “\32bit\Drivers\ USB3\win7\cyusb3.cat” ( c:\windows\inf\

- Copy “\32bit\Drivers\ USB3\win7\*.spt” ( c:\windows\system32\drivers\bwtek\

- Copy “\32bit\Drivers\USB3\win7\CyUSB3.sys” ( c:\windows\system32\Drivers\

- Copy “\32bit\Drivers\USB3\win7\WdfCoInstaller01009.dll” ( c:\windows\system32\

- Copy “\32bit\Drivers\USB3\win7\WdfCoInstaller01009.dll” ( c:\windows\system32\Drivers\

- Copy “\32bit\Drivers\CyUSB.dll” ( c:\windows\system32\

- Copy “\32bit\DLL_32bit\bwtekusb.dll” ( c:\windows\system32\

- Copy “\32bit\DLL_32bit\bwtekusb.dll” ( the folder of your executable file.

    - Windows 7 64-Bit

      - Windows 7 – 64-Bit - USB 2.0 Capable Units:

- Copy “\64bit \Drivers\USB2\win7\bwtekusb2.inf” ( c:\windows\inf\

- Copy “\64bit \Drivers\USB2\ win7\ fx2lp.cat” ( c:\windows\inf\

- Copy “\64bit \Drivers\USB2\ win7\*.spt” ( c:\windows\system32\drivers\bwtek\

- Copy “\64bit \Drivers\USB2\ win7\CyUSB.sys” ( c:\windows\system32\Drivers\

- Copy “\64bit \Drivers\CyUSB.dll” ( c:\windows\system32\

- Copy “\64bit\DLL_64bit\bwtekusb.dll” ( c:\windows\system32\

- Copy “\64bit\DLL_64bit\bwtekusb.dll” ( the folder of your executable file.

      - Windows 7 – 64-Bit - USB 3.0 Capable Units:

- Copy “\64bit\Drivers\USB3\win7\bwtekusb3.inf” ( c:\windows\inf\

- Copy “\64bit\Drivers\ USB3\win7\cyusb3.cat” ( c:\windows\inf\

- Copy “\64bit\Drivers\ USB3\win7\*.spt” ( c:\windows\system32\drivers\bwtek\

- Copy “\64bit\Drivers\USB3\win7\CyUSB3.sys” ( c:\windows\system32\Drivers\

- Copy “\64bit\Drivers\USB3\win7\WdfCoInstaller01009.dll” ( c:\windows\system32\

- Copy “\64bit\Drivers\USB3\win7\WdfCoInstaller01009.dll” ( c:\windows\system32\Drivers\

- Copy “\64bit\Drivers\CyUSB.dll” files ( c:\windows\system32\

- Copy “\64bit\DLL_64bit\bwtekusb.dll” ( c:\windows\system32\

- Copy “\64bit\DLL_64bit\bwtekusb.dll” ( the folder of your executable file.

  - Windows 8

    - Windows 8 - 32-Bit

      - Windows 8 - 32-Bit – USB 2.0 Capable Units:

- Copy “\32bit\Drivers\USB2\win8\bwtekusb2.inf” ( c:\windows\inf\

- Copy “\32bit\Drivers\USB2\win8\fx2lp.cat” ( c:\windows\inf\

- Copy “\32bit\Drivers\USB2\win8\*.spt” ( c:\windows\system32\drivers\bwtek\

- Copy “\32bit\Drivers\USB2\win8\CyUSB.sys” ( c:\windows\system32\Drivers\

- Copy “\32bit\Drivers\CyUSB.dll” ( c:\windows\system32\

- Copy “\32bit\DLL_32bit\bwtekusb.dll” ( c:\windows\system32\

- Copy “\32bit\DLL_32bit\bwtekusb.dll” ( the folder of your executable file.

      - Windows 8 - 32-Bit – USB 3.0 Capable Units:

- Copy “\32bit\Drivers\USB3\win8\bwtekusb3.inf” ( c:\windows\inf\

- Copy “\32bit\Drivers\USB3\win8\cyusb3.cat” ( c:\windows\inf\

- Copy “\32bit\Drivers\USB3\win8\*.spt” ( c:\windows\system32\drivers\bwtek\

- Copy “\32bit\Drivers\USB3\win8\CyUSB3.sys” ( c:\windows\system32\Drivers\

- Copy “\32bit\Drivers\USB3\win8\WdfCoInstaller01009.dll” ( c:\windows\system32\

- Copy “\32bit\Drivers\USB3\win8\WdfCoInstaller01009.dll” ( c:\windows\system32\Drivers\

- Copy “\32bit\Drivers\CyUSB.dll” files ( c:\windows\system32\

- Copy “\32bit\DLL_32bit\bwtekusb.dll” ( c:\windows\system32\

- Copy “\32bit\DLL_32bit\bwtekusb.dll” ( the folder of your executable file.

    - Windows 8 - 64-Bit

      - Windows 8 - 64-Bit – USB 2.0 Capable Units:

- Copy “\64bit\Drivers\USB2\win8\bwtekusb2.inf” ( c:\windows\inf\

- Copy “\64bit\Drivers\USB2\win8\fx2lp.cat” ( c:\windows\inf\

- Copy “\64bit\Drivers\USB2\win8\*.spt” ( c:\windows\system32\drivers\bwtek\

- Copy “\64bit\Drivers\USB2\win8\CyUSB.sys” ( c:\windows\system32\Drivers\

- Copy “\64bit\Drivers\CyUSB.dll” ( c:\windows\system32\

- Copy “\64bit\DLL_64bit\bwtekusb.dll” ( c:\windows\system32\

- Copy “\64bit\DLL_64bit\bwtekusb.dll” ( the folder of your executable file.

      - Windows 8 - 64-Bit – USB 3.0 Capable Units:

- Copy “\64bit\Drivers\USB3\win8\bwtekusb3.inf” ( c:\windows\inf\

- Copy “\64bit\Drivers\USB3\win8\cyusb3.cat” ( c:\windows\inf\

- Copy “\64bit\Drivers\USB3\win8\*.spt” ( c:\windows\system32\drivers\bwtek\

- Copy “\64bit\Drivers\USB3\win8\CyUSB3.sys” ( c:\windows\system32\Drivers\

- Copy “\64bit\Drivers\USB3\win8\WdfCoInstaller01009.dll” ( c:\windows\system32\

- Copy “\64bit\Drivers\USB3\win8\WdfCoInstaller01009.dll” ( c:\windows\system32\Drivers\

- Copy “\64bit\Drivers\CyUSB.dll” ( c:\windows\system32\

- Copy “\64bit\DLL_64bit\bwtekusb.dll” ( c:\windows\system32\

- Copy “\64bit\DLL_64bit\bwtekusb.dll” ( the folder of your executable file.

# USB 3.0/2.0/1.1 Interface Spectrometers

Below are the basic functions for the USB interface spectrometers for writing custom software.

Further in the document you will find Supplementary functions and model specific functions.

# USB Basic Functions

## InitDevices

BOOL InitDevices

(

);

This function initializes space in memory by creating a USB device object for all connected USB spectrometers.

***This is the first function which needs to be called.

RETURN

If the function call is successful, a True value will be returned.

## bwtekSetupChannel

int bwtekSetupChannel

(

int nFlag,				// Flag Value should always be -1

char *nChannelStatus	//Pointer to 32 byte array

);

This function reads the value of for the 32 available spectrometer channel numbers.

A total of 32 spectrometers can be used at one time, each with its own unique channel number.

Valid Channel Numbers for this array will be 0 – 31

*Note: Channel Numbers with this version of the SDK are now dynamically assigned.

Static Channels numbers are no longer able to be assigned.

nFlag -1 is used to get the value of all 32 available channel numbers.  
(This will work even if you do Not have 32 spectrometers connected)

nChannelStatus is a pointer to a 32 byte memory array which saves the channel-number values for all 32 available spectrometers.

*Note: values in the array which are >=32 are Inactive channels.

RETURN

If the function call is successful it will return a positive integer.  Otherwise it will return a negative integer.

## bwtekReadEEPROMUSB

int bwtekReadEEPROMUSB

(

char *OutFileName,		// The filename in which data from EEPROM will be saved

int nChannel			// Channel number of spectrometer received from the ‘bwtekSetupChannel’ function

);

This function is used for retrieving data from the spectrometer’s EEPROM

OutFileName is used to specify the name of the output file which will be saved to the computer. The output file has a text ASCII format and the default folder is the current directory unless otherwise specified.

File Naming Examples:

bwtekReadEEPROMUSB(“C:\para.ini”,0)

bwtekReadEEPROMUSB(“eeprom.txt”,1)

nChannel is used to address a specific spectrometer device to be operated when multiple spectrometer devices are involved.  Users must call the bwtekSetupChannel function to determine which channel numbers are available.  A total of 32 spectrometer devices may be connected at one time, where the nChannel value will range from 0 – 31.

RETURN

FAIL: If there is No active spectrometer assigned to a given channel number that is being read a -1 value will be returned.

PASS: A Positive value be returned AND a output file will be created in the folder location designated from above.

NOTE:  The output file contains a lot of information used for internal use.  Some of this data, for example, the C-code for the spectrometer, calibration coefficients, timing mode, input mode and number of pixels on the spectrometer’s detector may be useful.

## bwtekTestUSB

int bwtekTestUSB

(

int nTimingMode,	// USB Interface timing option
   int nPixelNo,		// number of pixels of a detector to be readout 
   int nInputMode,	// signal conditioning stage gain value
   int nChannel,		// channel to get data from 
   int pParam 		// setting for RS232 – Use NULL FOR USB INTERFACES
);

This function is for initializing communication to a specific USB spectrometer device while knowing its specific nChannel number assignment.

nTimingMode is used to specify the USB firmware timing option.

Note:  See Appendix E for nTimingMode values for the spectrometer model you are using OR this information can be retrieved from the *OutFileName created from the bwtekReadEEPROMUSB function.

nPixelNo should be set to the number of pixels of used by the detector array in the spectrometer device.

Note:  See Appendix E for nPixelNo values for the spectrometer model you are using OR this information can be retrieved from the *OutFileName created from the bwtekReadEEPROMUSB function.

nInputMode is used to specify the ADC input range used in the spectrometer device being programmed. It should be set to 0 for ADCs for unipolar input such as 0 to +5V or 0 to +10V. It should be set to 1 for ADCs using bipolar input such as -5 to +5V or -10 to +10V (BTC and BRC111), and 2 for ADCs for unipolar and non-inverting input (BTC600).

Note:  See Appendix E for nInputMode values for the spectrometer model you are using OR this information can be retrieved from the *OutFileName created from the bwtekReadEEPROMUSB function.

nChannel is used to address a specific spectrometer device to be operated when multiple spectrometer devices are involved.  Users must call the bwtekSetupChannel function to determine which channel numbers are available.  A total of 32 spectrometer devices may be connected at one time, where the nChannel value will range from 0 – 31.

pParam This pointer should be NULL FOR ALL USB INTERFACES

Return

If the spectrometer readout is successful, a positive integer or 0 will be returned.

## bwtekSetTimeUSB

long bwtekSetTimeUSB

(

long lTime, 	// integration time setting
  int nChannel 	// channel to get data from
);

This function is for setting the spectrometer integration time as specified.

*BRC115E, BRC115U, and BRC115V model, lTime always is microsecond.

lTime Range: 1,000 ( 2,100,000,000 (us)

*BTC261P, BTC264P model, lTime always is microsecond.

lTime Range: 200 ( 2,100,000,000 (us)

*BTC655E, BTC665E model, lTime always is microsecond.

lTime Range: 6,000 ( 2,100,000,000 (us)

lTime is the integration time value to be set

*Used in conjunction with the bwtekSetTimeUnitUSB function above.

Refer to Appendix F: Default Parameters for bwtekTESTUSB for integration time range for your spectrometers.

(1) For spectrometers that have 2D detectors installed, (BTC6xx series), an offset integration time, bwtekSetTimeBase0USB, results from the extra time the sensing pixels are exposed to during the readout process of the area sensor arrays.  The actual Exposure time of the sensor is, set integration time + offset time.  The bwtekSetTimeBase0USB needs to be taken into account when calculating the actual integration time.  Therefore the lTime value to be passed to the DLL is calculated by subtracting the offset time from the desired integration time.

For Example:

If your Offset Time is 26 (ms) and desired integration time is 50 (ms).

You would NOT set lTime to 50

You MUST set lTime = (desired exposure time) – Offset Time

lTime = 50 – 26 --> lTime = 24

This will cause the minimal integration time for the spectrometer to be affected. The minimal integration time will be Offset Time + 1.

(2) For the BTC263 model, if your integration time is set in microseconds (us), the actual exposure time you will receive is your desired exposure time + 80us.

For Example:

If you want to set the integration time to 200 (us) [desired exposure time]

You MUST set lTime = (desired exposure time) + Offset Time (80 us)

lTime = 200 + 80 --> lTime = 280 (us)

If the BTC263 model integration time is set to 1 millisecond (ms), there is NO Offset Time.

For Example: 	If you want to set the integration time to 10 (ms) [desired exposure time]

lTime = 10 (ms)

nChannel is used to address a specific spectrometer device to be operated when multiple spectrometer devices are involved.  Users must call the bwtekSetupChannel function to determine which channel numbers are available.  A total of 32 spectrometer devices may be connected at one time, where the nChannel value will range from 0 – 31.

RETURN

If the function call is successful it returns the new integration time value. Otherwise it will return a negative value.

## bwtekSetTimeBase0USB

long bwtekSetTimeBase0USB

(

long lOffsetTime, 	// offset integration time setting
  int nChannel 		// channel to get data from
);

This function is for setting the spectrometer offset integration time as specified.

lOffsetTime is the offset time base value in milliseconds to be set

.

Refer to Appendix B for Offset Integration Time settings for your spectrometer.

*If the Offset Integration Time value in Appendix B is ‘NA’ you do Not need to call this function for your spectrometer.

nChannel is used to address a specific spectrometer device to be operated when multiple spectrometer devices are involved.  Users must call the bwtekSetupChannel function to determine which channel numbers are available.  A total of 32 spectrometer devices may be connected at one time, where the nChannel value will range from 0 – 31.

RETURN

If the function call is successful it returns the new offset integration time value. Otherwise it will return negative value.

## bwtekSetTimingsUSB

int bwtekSetTimingsUSB

(

long lTriggerExit,		// Setting external trigger timeout

int nMultiple, 			// A multiplier factor for integration times needed > 65,535 ms

int nChannel			//channel number to get data from

);

This function is used for setting both an external trigger timeout and a multiplier factor when an integration time is needed > 65,535ms.

lTriggerExit is used to specify the wait time for receipt of an external trigger. If no trigger signal is received within this set time, the spectrometer will automatically take a single scan.  
The real timeout period is 15ms * lTriggerExit,

Maximum timeout period: 15ms * 65,535 = 983,040ms.

**To avoid the spectrometer/software from timing out and continuously wait for an external trigger, the lTriggerExit value should equal 0

nMultiple is used to specify the multiplying factor for the integration time.  The default time base is 1 when multiple = 1, when multiple = 2, the time base is 2 and real integration time will be 2 * lTime (lTime parameter is from the bwtekSetTimeUSB function. The lTime time can be adjusted by using bwtekSetTimeUSB function.  The real integration time period is [nMultiple * ( lTime, - lOffsetTime )]

nMultiple range is 1 - 16.

nChannel is used to address a specific spectrometer device to be operated when multiple spectrometer devices are involved.  Users must call the bwtekSetupChannel function to determine which channel numbers are available.  A total of 32 spectrometer devices may be connected at one time, where the nChannel value will range from 0 – 31.

RETURN

If the spectrometer device becomes disconnected or loses communication a negative integer will be returned, indicating a failure.  Otherwise the lTriggerExit value will be returned.

## bwtekDataReadUSB

int bwtekDataReadUSB

(

int nTriggerMode,			// Base address for plug-in data acquisition board 
  unsigned short *pArray, 		// data value array from the read operation stored

int nChannel 				// channel to get data from

);

This function is for reading out data form detector

nTriggerMode is used to set the trigger mode to initiate a trigger scan process. It should be set to 0 for free running (continuous scanning) mode and 1 for external trigger mode.

The external trigger signal should be supplied as a 5V TTL tpulse.

It is falling edge effective.

*Refer to the spectrometer’s hardware user manual for the trigger pulse width definition.

pArray is a pointer to a data array memory space, size will depend on the total number of pixels on the CCD for the spectrometer model that is being used.  Every element in this array should be an unsigned integer with a minimum of 2 bytes for 16 bit resolution spectrometers

.
NOTE:  The pArray size should be equal to int nPixelNo in the bwtekTestUSB function.

See Appendix A for Pixel Number Values

nChannel is used to address a specific spectrometer device to be operated when multiple spectrometer devices are involved.  Users must call the bwtekSetupChannel function to determine which channel numbers are available.  A total of 32 spectrometer devices may be connected at one time, where the nChannel value will range from 0 – 31.

RETURN

If the spectrometer readout is successful the number of data points read will be returned.

Else a negative will be returned.

Notes:

The pArray buffer needs to be REVERSED if the EEPROM contains the “xaxis_data_reverse=1” entry.

Refer to the X-axis Reverse Appendix for the setting relating to specific model types.

pArray[0] (pArray[nPixelNo-1]

pArray[1] (pArray[nPixelNo-2]

.....

pArray[nPixelNo-2] (pArray[1]

pArray[nPixelNo-1] (pArray[0]

Example how to get the “xaxis_data_reverse=1” entry from the EEPROM:

(1) Read EEPROM context to para.ini file (bwtekReadEEPROMUSB(“para.ini”,0)

(2) Search for x-axis reverse flag in the [COMMON] section of the para.ini file:

[COMMON]

xaxis_data_reverse=1

By default the xaxis_data_reverse is set to 0.

## bwtekReadResultUSB

int bwtekReadResultUSB

(

int nTriggerMode, 			// setting for trigger mode

int nAverage, 				// number of scans to be averaged

int nTypeSmoothing,		// setting for smoothing type

int nValueSmoothing, 		// setting for smoothing parameter

unsigned short *pArray,	 	// an array memory space for scan data

int nChannel		 		// setting for channel number

);

This function is for reading out data from the detector then applying data smoothing.

nTriggerMode is used to set the trigger mode to initiate a trigger scan process. It should be set to 0 for free running (continuous scanning) mode and 1 for external trigger mode. The external trigger signal should be supplied as a 5V TTL tpulse. It is falling edge effective.

Refer to hardware user manual for the trigger pulse width definition.

nAverage is used to set the number of scans to be averaged.  (Set value to 1 for no averaging)

nAverage range 1 - 16

nTypeSmoothing 0 for no smoothing function, 1 for FFT smoothing, 2 for Savitzky-Golay smoothing, 3 for Boxcar smoothing.

nValueSmoothing

- When using FFT smoothing (nTypeSmoothing=1), this parameter indicates the percentage of cutoff frequency.  The nValueSmoothing should be 0 to 100.

- When using Savitzky-Golay smoothing (nTypeSmoothing=2), The nValueSmoothing should be 2 to 5.

Example:

When using Savitzky-Golay smoothing, if nValueSmoothing is set to 4, the total number of smoothing points will be 9.  1 point will be the “origin” pixel and the other 8 will be 4 pixels before the origin and 4 pixels after the origin added together with weighted factors and divided by 9.

- When using Boxcar smoothing (nTypeSmoothing=3), The nValueSmoothing should be the number of pixels to smooth.

pArray is a pointer to the data array’s memory space dependent on the total number of pixels on the CCD for the spectrometer model that is being used.  Every element in this array should be an unsigned integer with minimum 2 bytes for a 16 bit resolution.

NOTE:  This Value should be equal to int nPixelNo in the bwtekTestUSB function.

See Appendix A for Pixel Number Values

nChannel is used to address a specific spectrometer device to be operated when multiple spectrometer devices are involved.  Users must call the bwtekSetupChannel function to determine which channel numbers are available.  A total of 32 spectrometer devices may be connected at one time, where the nChannel value will range from 0 – 31.

RETURN

If the spectrometer readout is successful as the result of this function call will be 0.  Otherwise the return value is negative.

Notes:

The pArray buffer needs to be REVERSED if the EEPROM contains the “xaxis_data_reverse=1” entry.

Refer to Appendix F X-axis Inverse Settings.

pArray[0] (pArray[nPixelNo-1]

pArray[1] (pArray[nPixelNo-2]

.....

pArray[nPixelNo-2] (pArray[1]

pArray[nPixelNo-1] (pArray[0]

Example how to get the “xaxis_data_reverse=1” entry from the EEPROM:

(1) Read EEPROM context to para.ini file (bwtekReadEEPROMUSB(“para.ini”,0)

(2) Search for x-axis reverse flag in the [COMMON] section of the para.ini file:

[COMMON]

xaxis_data_reverse=1

By default the xaxis_data_reverse is set to 0.

## bwtekDataReadUSB1

int bwtekDataReadUSB1

(

int nScanNo           		// scan number

int nTriggerMode,			// base address for plug-in data acquisition board 
  unsigned short *pArray, 		// data value array from the read operation stored

double *pStartTime,      	// array save the start time of scan data, unit is ms

double *pEndTime,       	// array save the end time of scan, unit is ms

int nChannel 				// channel to get data from

);

This function is for reading out data from detector

nScanNo is the number of data scans.

nTriggerMode is used to set the trigger mode to initiate a trigger scan process. It should be set to 0 for free running (continuous scanning) mode and 1 for external trigger mode. The external trigger signal should be supplied as a 5V TTL tpulse. It is falling edge effective.

Refer to hardware user manual for the trigger pulse width definition.

pArray is a pointer to the data array’s memory space dependent on the total number of pixels on the CCD for the spectrometer model that is being used.  Every element in this array should be an unsigned integer with minimum 2 bytes for a 16 bit resolution.

NOTE:  This Value should be equal to int nPixelNo in the bwtekTestUSB function.

See Appendix A for Pixel Number Values

pStartTime is a pointer to a start time array’s memory space dependent on the number of data scans  nScanNo.

pEndTime is a pointer to an end time array’s memory spaces dependent on the number of data scans  nScanNo.

nChannel is used to address a specific spectrometer device to be operated when multiple spectrometer devices are involved.  Users must call the bwtekSetupChannel function to determine which channel numbers are available.  A total of 32 spectrometer devices may be connected at one time, where the nChannel value will range from 0 – 31.

RETURN

If the spectrometer readout is successful as the result of this function call it returns the positive value. Otherwise the return value is negative.

Notes:

The pArray buffer needs to be REVERSED if the EEPROM contains the “xaxis_data_reverse=1” entry.

Refer to Appendix F X-axis Inverse Settings.

pArray[0] (pArray[nPixelNo-1]

pArray[1] (pArray[nPixelNo-2]

.....

pArray[nPixelNo-2] (pArray[1]

pArray[nPixelNo-1] (pArray[0]

Example how to get the “xaxis_data_reverse=1” entry from the EEPROM:

(1) Read EEPROM context to para.ini file (bwtekReadEEPROMUSB(“para.ini”,0)

(2) Search for x-axis reverse flag in the [COMMON] section of the para.ini file:

[COMMON]

xaxis_data_reverse=1

By default the xaxis_data_reverse is set to 0.

## bwtekGetTimeUSB

long bwtekGetTimeUSB

(

long *lTime, 		// reads back the integration time setting
  int nChannel 		// channel to get data from
);

This function is for reading back the spectrometer’s set integration time (lTime) which was set in the bwtekSetTimeUSB function.

*lTime is a pointer to an array that will return the integration time value (lTime) which was set in the bwtekSetTimeUSB function.

nChannel is used to address a specific spectrometer device to be operated when multiple spectrometer devices are involved.  Users must call the bwtekSetupChannel function to determine which channel numbers are available.  A total of 32 spectrometer devices may be connected at one time, where the nChannel value will range from 0 – 31.

RETURN

If the function call is successful it returns the value 1. Otherwise it will return negative value or lTime value equal 0.

## GetDeviceCount

int GetDeviceCount

(

);

This function is returns the number of spectrometer connected to the host computer.

RETURN

If the function call is successful, a positive integer will be returned, else a negative integer will be returned.

## GetUSBType

int GetUSBType

(

int *nUSBType,	 	// pointer to an array for the spectrometer’s USB Type

int nChannel		// channel to get data from

);

This function will return the spectrometer’s USB type.

*nUSBType is a pointer to an array that will return the physical USB Type interface of the connected spectrometer unit.

nUSBType = 2, it is USB2.0 unit,  nUSBType=3, it is USB3.0 unit.

nChannel is used to address a specific spectrometer device to be operated when multiple spectrometer devices are involved.  Users must call the bwtekSetupChannel function to determine which channel numbers are available.  A total of 32 spectrometer devices may be connected at one time, where the nChannel value will range from 0 – 31.

RETURN

If the function call is successful it returns the value 1. Otherwise it will return a negative value.

## bwtekGetCCode (Obsolete)

OBSOLETE - Use GetCC0de function below.

## GetCCode

int GetCCode

(

byte *CCode,		// returned CCode of spectrometer

int nChannel 		// channel to get data from

);

This function is for reading the c-code of the spectrometer off the spectrometer’s EEPROM.

*CCode is a pointer to a 32-byte array which contains the c-code of the spectrometer.

nChannel is used to address a specific spectrometer device to be operated when multiple spectrometer devices are involved.  Users must call the bwtekSetupChannel function to determine which channel numbers are available.  A total of 32 spectrometer devices may be connected at one time, where the nChannel value will range from 0 – 31.

RETURN

If the function call is successful, a positive integer will be returned, else a negative integer will be returned.

## bwtekCloseUSB

int bwtekCloseUSB

(

int nChannel 	// channel to get data from

);

This function is for ending communication to a specific USB spectrometer device while knowing its specific nChannel number assignment.  This should be used when closing the programmed spectrometer application which was initiated by the bwtekTestUSB function call.

nChannel is used to address a specific spectrometer device to be operated when multiple spectrometer devices are involved.  Users must call the bwtekSetupChannel function to determine which channel numbers are available.  A total of 32 spectrometer devices may be connected at one time, where the nChannel value will range from 0 – 31.

RETURN

If closing USB connection is successful as the result of this function call it returns positive integer otherwise the return value is negative.

## CloseDevices

int CloseDevices

(

);

This function is for closing all USB Devices.  This is the last function that should be called when closing your software program which was initialed by the InitDevices function call.

RETURN

If the function call is successful, a positive integer will be returned, else a negative integer will be returned.

# Supplementary Functions

## bwtekSmoothingUSB

int bwtekSmoothingUSB

(

int nTypeSmoothing,			// Smoothing type setting

int nValueSmoothing, 		// Smoothing parameter

double *pArray, 	    			// an array memory space for desire smoothing data

int nNum     		    		// array length

);

This function is for applying data smoothing..

nTypeSmoothing 0 for BoxCar , 1 for FFT smoothing, 2 for Savitzky-Golay smoothing

nValueSmoothing

- BoxCar smoothing (nTypeSmoothing = 0), nValueSmoothing = number of pixels to smooth.

- FFT smoothing (nTypeSmoothing = 1), nValueSmoothing should be 0 to 100.  This indicates the percentage of cutoff frequency.

- Savitzky-Golay smoothing (nTypeSmoothing = 2), nValueSmoothing should be 2 to 5.

Example:

When using Savitzky-Golay smoothing, if nValueSmoothing is set to 4, the total number of smoothing points will be 9.  1 point will be the “origin” pixel and the other 8 will be 4 pixels before the origin & 4 pixels after the origin added together with weighted factors and divided by 9.

pArray is a pointer to an array containing the desired smoothing variables. It should range from 0 to (nNum-1).   After the bwtekSmoothingUSB function is finished, this array will contain the smoothed data.

nNum is the total number of data points (length) in pArray

RETURN

If the function call is successful it will return a positive integer.  Otherwise it will return a negative integer.

Note: Setting nTypeSmoothing = 0 and nTypeSmoothing = 0, will perform no smoothing.

## bwtekConvertDerivativeDouble

int bwtekConvertDerivativeDouble

(

int nSelectMethod,

int nPointHalf,

int nPolynominalOrder,

int nDerivativeOrder,

double *pRawArray,

double *pResultArray,

int nNum

);

This function is for data derivative.

nSelectMethod 0 for Point Diff, 1 for Savitzky-Golay.

nPointHalf is half of selected points.

nPolynominalOrder is degree of polynomial.

nDerivativeOrder is order of derivate. 0 is 1st derivate. 1 is 2nd derivate.

pRawArray is pointer of desired derivate data array.

pResultArray is pointer of derivate data array.

nNum is the total number of data points (length) in pRawArray

RETURN

If the function call is successful it will return a positive integer.  Otherwise it will return a negative integer.

## bwtekPolyFit

int bwtekPolyFit

(

double *x,			// Array of independent variables

double *y, 			// Array of dependent variables

int const numPts, 	// Number of points in the independent and dependent arrays

double *coefs, 		// Pointer to array containing calculated coefficients [index from 0 to ‘order’ parameter]

int const order	 	// Desired order of polynomial fit

);

This function is to perform a curve fit to a specified polynomial function based on supplied data values and ‘least square’ algorithm.  This function may be used for converting pixel numbers to wavelengths.

If used for wavelength conversion, a wavelength calibration step needs to be performed, during which a calibration light source with known spectral features or a series of narrow band light radiation will be necessary.  The known wavelength features of the calibration source along with the pixel numbers they are falling onto are the Y and X data values that need to be passed to the function for curve fitting.  The array indices should range from 0 -- [Number_of_Points - 1]. They are passed to the function along with a desired polynomial fit order and an array large enough to hold the coefficients ().  This array may be used with bwtekPolyCalc to calculate wavelength from pixels.

x is a pointer to an array containing the independent variables. It should range from 0 to (numPts-1).

y is a pointer to an array containing the dependent variables. It should range from 0 to (numPts-1).

numPts is the total number of data pairs in the variable arrays

coefs is a pointer to an array which contains the polynomial curve fit coefficients whose number of memory spaces are (order +1).

order is the desired polynomial curve fit order.  The third order fit is recommended for use for pixel to wavelength conversions.

RETURN

If the function call is successful it will return a positive integer.  Otherwise it will return a negative integer.

## bwtekPolyCalc

void bwtekPolyCalc

(

double *coefs,		// pointer to the polynomial coefficients

int const order,		// polynomial order to use

double const x,		// pixel number

double *y			// wavelength value

);

This function performs calculations by using the following formula:

y = a0 + a1*x^1 + a2*x^2 + ... + aN*x^N

coefs is a pointer to an array containing the polynomial coefficients. These can be calculated using the bwtekPolyFit function from above.

order is the polynomial order to be used and must be less than or equal to the length of the coefs array.

x is the input variable, in this case, the pixel number.

y is the value to be calculated, in this case, wavelength (nm).

RETURN

If the function call is successful it will return a positive integer.  Otherwise it will return a negative integer.

## bwtekDataExport

int bwtekDataExport

(

DataExport_Parameter_Struct *pDataExport_Parameter_Struct,	// parameter structure of export

SpectrometerCoefs_Struct *pSpectrometerCoefs_Struct,  		// parameter structure of spectrometer coeff.

double *pSrcArray,                                   		// pointer to pixel data array

double *pResultArray,                    					// pointer to converted data

int *ResultArrayLength									// length of ‘ResultArray’

);

This function is to be used to export x-axis data evenly.

X-axis units that can be exported evenly are the following: Wavelength / Wavenumber / Raman shift.

DataExport_Parameter_Struct is structure of export parameters.

typedef struct DataExport_Parameter_Struct

{

int nOptionX;			// 0 for wavelength, 1 for Wavenumber, 2 for Raman shift

int nStartX;       	// x-axis start location based on format (nOptionX)

int nEndX;			// x-axis end location based on format (nOptionX)

int nIncX;          	// x-axis increments based on format (nOptionX)

int nPixelNumber;		// Spectrometer detector’s total number of Pixels (Refer to Appendix B)

}

SpectrometerCoefs_Struct is the structure of the spectrometer’s coefficients.

typedef struct SpectrometerCoefs_Struct

{

double fExcitationWavelength; 	// FOR RAMAN SHIFT ONLY – Measured Excitation wavelength of Laser

double fCoefsA[4];           	// coefficients for pixel conversion to wavelength

double fCoefsB[4]; 			// coefficients for wavelength conversion to pixel

}

- fExcitationWavelength is the value of measured laser excitation wavelength for Raman shift

- fCoefsA are the A coefficients for pixel to wavelength conversion.  Can be found on the EEPROM of the spectrometer, “coefs_a0... coefs_a3”

- fCoefsB are the B coefficients for wavelength to pixel conversion.  Can be found on the EEPROM of the spectrometer, “coefs_b0... coefs_b3”

pSrcArray is a pointer to the data source array, the data to be converted.  This data will be in pixel format.

pResultArray is a pointer to the data which has been converted from the data source.  This data will be in either evenly spaced; Wavelength, Wavenumber or Raman shift values.

pResultArrayLength is the length of the pResultArray.

RETURN

If the function call is successful it will return a positive integer.  Otherwise it will return a negative integer.

Notes:

How to set the parameter of this function:

Example:

BTC112 Spectrometer coefficient’s are as follows:

Coefs_a0…a3=321.107087610289 / 0.446657617771052 / -3.40125796972623E-5 / -1.27782145731415E-9

Coefs_b0…b3= -728.75518074818 / 2.29668100409253 / -0.00025083531593495 / 4.59557309079057E-7

Real Excitation wavelength= 784.85 nm

|  | Wavelength (Evenly) 600--900 nm | Wave Number (Evenly) 19000--11000 (cm-1) | Raman Shift (Evenly) -6500 -- 1500 (cm-1) |
| --- | --- | --- | --- |
| nOptionX | 0 | 1 | 2 |
| nStartX | 600 (nm) | 19000 (cm-1) | -6500 (cm-1) |
| nEndX | 900 (nm) | 11000 (cm-1) | 1500 (cm-1) |
| nIncX | 1 (nm) | 1 (cm-1) | 1 (cm-1) |
| nPixelNumber | 2048 | 2048 | 2048 |
| fExcitationWavelength | 784.85 (nm) | 784.85 (nm) | 784.85 (nm) |
| fCoefsA [0] | 321.107087610289 | 321.107087610289 | 321.107087610289 |
| fcoefsA [1] | 0.446657617771052 | 0.446657617771052 | 0.446657617771052 |
| fCoefsA [2] | -3.40125796972623E-5 | -3.40125796972623E-5 | -3.40125796972623E-5 |
| fCoefsA [3] | -1.27782145731415E-9 | -1.27782145731415E-9 | -1.27782145731415E-9 |
| fCoefsB [0] | -728.75518074818 | -728.75518074818 | -728.75518074818 |
| fcoefsB [1] | 2.29668100409253 | 2.29668100409253 | 2.29668100409253 |
| fcoefsB [2] | -0.00025083531593495 | -0.00025083531593495 | -0.00025083531593495 |
| fcoefsB [3] | 4.59557309079057E-7 | 4.59557309079057E-7 | 4.59557309079057E-7 |

## bwtekSaveEEPROMChannel (Obsolete)

OBSOLETE – Channel Numbers are now dynamically assigned.

## bwtekGetXaxisInverseByte

int bwtekGetXaxisInverseByte

(

int *nXaxisInverseByte   	// returned x-axis inverse flag of the spectrometer

int nChannel 		 	// channel to get data from

);

This function is for reading the x-axis inverse flag of the spectrometer from the spectrometer’s EEPROM.

*nXaxisInverseByte is a pointer to an array that will contain the x-axis inverse value of the spectrometer.

(1 = x-axis inversed, 0 = x-axis NOT inversed)

nChannel is used to address a specific spectrometer device to be operated when multiple spectrometer devices are involved.  Users must call the bwtekSetupChannel function to determine which channel numbers are available.  A total of 32 spectrometer devices may be connected at one time, where the nChannel value will range from 0 – 31.

RETURN

If the function call is successful, a positive integer will be returned, else a negative integer will be returned.

# Temperature Readout – Limited Use

Reference Temperature Sensing Appendix for a list of spectrometer which support the use of this function call.

## bwtekReadTemperature

int bwtekReadTemperature

(

int nCommand,

int *nADValue,			\\ AD Value Range: 0 - 4095

double *nTemperature, 	\\ Temperature will be in Celsius

int nChannel

);

This function is for reading the temperature of the CCD Detector OR read the ambient temperature.

nCommand is a flag of specific temperature: 
0x10 for CCD detector temperature, 
0x11 for ambient temperature reading.

*nADValue is a pointer to an integer return value. It is reserved.

*nTemperature is a pointer to float pointer value of temperature. The temperature units are degrees Celsius.

nChannel is used to address a specific spectrometer device to be operated when multiple spectrometer devices are involved.  Users must call the bwtekSetupChannel function to determine which channel numbers are available.  A total of 32 spectrometer devices may be connected at one time, where the nChannel value will range from 0 – 31.

RETURN

If the function call is successful, a positive integer will be returned, else a negative integer will be returned.

# USB 3.0 Interface Spectrometers

Only spectrometer with a USB 3.0 port on the spectrometer can use the functions listed in this section, ‘USB 3.0 Interface Spectrometers’

## bwtekDSPDataReadUSB

int bwtekDSPDataReadUSB

(

int nAveNum,

int nSmoothing,

int nDarkCompensate,

int nTriggerMode,

unsigned short *pArray,

int nChannel

);

‘Smart Scan Mode’

This function applies onboard averaging, smoothing, dark compensate on the firmware side of the spectrometer before outputting the data to the computer.

*Note:  This function will work ONLY with spectrometer with a USB 3.0 interface connector.

nAveNum is used to set the number of scans to be averaged, (Range 0 – 65535)

nSmoothing is used to apply smoothing to the scanned data.  0 for no smoothing function, 1 for Savitzky-Golay smoothing.

*NOTE: Only Savitzky-Golay is available for smoothing.

nDarkCompensate is used to apply dark compensate to the scanned data.  0 for no dark compensate function, 1 for dark compensate function.

nTriggerMode is used to set trigger mode to initiate a trigger scan process.

0 = Free Running Mode, 1 = External Trigger Mode.  
The external trigger signal should be supplied as a 5V TTL pulse.  It is falling edge effective.  
Refer to hardware user manual for the trigger pulse width definitions.

pArray is a pointer to data array memory spaces dependent on the total number of pixels on the CCD for the spectrometer model that is being used.  Every element in this array should be an unsigned integer with minimum 2 bytes for a 16 bit resolution.
*NOTE:  This Value should be equal to int nPixelNo in the bwtekTestUSB function.

nChannel is used to address a specific spectrometer device to be operated when multiple spectrometer devices are involved.  Users must call the bwtekSetupChannel function to determine which channel numbers are available.  A total of 32 spectrometer devices may be connected at one time, where the nChannel value will range from 0 – 31.

RETURN

If the spectrometer readout is successful the number of data points read will be returned. Otherwise a negative value will be returned.

## bwtekFrameDataReadUSB

int bwtekFrameDataReadUSB

(

int nFrameNum,

int nTriggerMode,

unsigned short *pArray,

int nChannel

);

*Update Note: This Function was formally called bwtekDataReadUSB

‘Burst Mode’

This function allows for the spectrometer to acquire X number of scans (nFrameNum) onboard on the firmware side of the spectrometer before outputting the data to the computer.

*Note:  This function will work ONLY with spectrometer with a USB 3.0 interface connector.
For Example:  BRC115E

.

nFrameNum is the number of scans to be acquired.

nTriggerMode is used to set trigger mode to initiate a trigger scan process.

0 = Free Running Mode, 1 = External Trigger Mode.  
The external trigger signal should be supplied as a 5V TTL pulse.  It is falling edge effective.  
Refer to hardware user manual for the trigger pulse width definitions.

pArray is a pointer to data array memory spaces dependent on the total number of pixels on the CCD for the spectrometer model that is being used.  Every element in this array should be an unsigned integer with minimum 2 bytes for a 16 bit resolution.
*NOTE:  This Value should be equal to int nPixelNo in the bwtekTestUSB function.

nChannel is used to address a specific spectrometer device to be operated when multiple spectrometer devices are involved.  Users must call the bwtekSetupChannel function to determine which channel numbers are available.  A total of 32 spectrometer devices may be connected at one time, where the nChannel value will range from 0 – 31.

RETURN

If the spectrometer readout is successful the number of data points read will be returned. Otherwise a negative value will be returned.

## bwtekWriteBlockUSB

int bwtekWriteBlockUSB

(

unsigned int nAddress,

byte *pArray,

int nNum,

int nChannel

);

This function is to write to a reserved location on the EEPROM by the customer.

This memory block must be first erased before rewriting to this location.

See bwtekEraseBlockUSB function below.

*NOTE: Only USB3.0 unit support this function.

nAddress is an address in EEPROM to write, location: 0 – 0xFFFF bytes.

pArray points to the data array memory space.

nNum is the number of desired bytes to be written: (0 – 0xFFFF bytes)
*NOTE: nAddress + nNum MUST BE <= 0xFFFF

nChannel is used to address a specific spectrometer device to be operated when multiple spectrometer devices are involved.  Users must call the bwtekSetupChannel function to determine which channel numbers are available.  A total of 32 spectrometer devices may be connected at one time, where the nChannel value will range from 0 – 31.

RETURN

If the function call is successful, a positive integer will be returned, else a negative integer will be returned

## bwtekReadBlockUSB

int bwtekReadBlockUSB

(

unsigned int nAddress,

byte *pArray,

int nNum,

int nChannel

);

This function is to Read the reserved location on the EEPROM written to by the customer.

*NOTE: Only USB3.0 unit support this function.

nAddress is an address in EEPROM to read, location: 0 – 0xFFFF bytes.

pArray points to the data array memory space.

nNum is the number of desired bytes to be read: (0 – 0xFFFF bytes)
*NOTE: nAddress + nNum MUST BE <= 0xFFFF

nChannel is used to address a specific spectrometer device to be operated when multiple spectrometer devices are involved.  Users must call the bwtekSetupChannel function to determine which channel numbers are available.  A total of 32 spectrometer devices may be connected at one time, where the nChannel value will range from 0 – 31.

RETURN

If the function call is successful, a positive integer will be returned, else a negative integer will be returned

## bwtekEraseBlockUSB

int bwtekEraseBlockUSB

(

int nChannel

);

This function is to ERASE the entire block memory reserved location (0 – 0xFFFF bytes) on the EEPROM written to by the customer.

*NOTE: Only USB3.0 unit support this function.

nChannel is used to address a specific spectrometer device to be operated when multiple spectrometer devices are involved.  Users must call the bwtekSetupChannel function to determine which channel numbers are available.  A total of 32 spectrometer devices may be connected at one time, where the nChannel value will range from 0 – 31.

RETURN

If the function call is successful, a positive integer will be returned, else a negative integer will be returned

# AUX Port Functions 1 - Multi-Purpose TTL Output Functions

For older spectrometers, in their user manual they may have different names for TTL Input, such as TTL Sensing.   The term TTL Sensing and TTL Input are one in the same.

**Reference the Appendices, SDK Function Groups, TTL Output and TTL Input to verify if your spectrometer model can use the below function calls.

## bwtekGetExtStatus

int bwtekGetExtStatus

(

int nChannel  	//channel to get data from

);

This function is used to retrieve the status of the TTL Input pin.

nChannel is used to address a specific spectrometer device to be operated when multiple spectrometer devices are involved.  Users must call the bwtekSetupChannel function to determine which channel numbers are available.  A total of 32 spectrometer devices may be connected at one time, where the nChannel value will range from 0 – 31.

RETURN

If the status of input pin is low, a 0 will be returned. If the status of input pin is high, a positive integer will be returned.  If the function call is unsuccessful, a negative integer will be returned.

## bwtekSetExtLaser

int bwtekSetExtLaser

(

int nOnOff, 	//switch for On/Off

int nChannel  	//channel to get data from

);

This function is used for controlling the On/Off capabilities of an external device, typically used for Laser control, through the Spectrometer’s AUX Port, when wired correctly.

nOnOff is used to set the external control signal. If 0 is assigned, the output pin signal will be low.  If 1 is assigned, the output pin signal will be high.

nChannel is used to address a specific spectrometer device to be operated when multiple spectrometer devices are involved.  Users must call the bwtekSetupChannel function to determine which channel numbers are available.  A total of 32 spectrometer devices may be connected at one time, where the nChannel value will range from 0 – 31.

RETURN

If the function call is successful, a positive integer will be returned, else a negative integer will be returned.

## bwtekSetExtShutter

int bwtekSetExtShutter

(

int nOnOff,		// switch for On/Off

int nChannel  	// channel to get data from

);

This function is used for controlling the Open/Closed capabilities of a Shutter (If Installed).

nOnOff is used to set the Shutter control signal. If 0 is assigned the output signal will be low.  If 1 is assigned the output signal will be high.

nChannel is used to address a specific spectrometer device to be operated when multiple spectrometer devices are involved.  Users must call the bwtekSetupChannel function to determine which channel numbers are available.  A total of 32 spectrometer devices may be connected at one time, where the nChannel value will range from 0 – 31.

RETURN

If the function call is successful, a positive integer will be returned, else a negative integer will be returned.

## bwtekSetExtSync

int bwtekSetExtSync

(

int nOnOff, 	//switch for On/Off

int nChannel  	//channel to get data from

);

This function is used to generate a pulse, which synchronizes an outside source with the external trigger signal for the spectrometer for detector scanning.

nOnOff is used to control the synchronizing pulse.  If 0 assigned, the output signal will not be generated (no signal). If 1 is assigned, the output signal will be generated (pulse signal).

nChannel is used to address a specific spectrometer device to be operated when multiple spectrometer devices are involved.  Users must call the bwtekSetupChannel function to determine which channel numbers are available.  A total of 32 spectrometer devices may be connected at one time, where the nChannel value will range from 0 – 31.

RETURN

If the function call is successful, a positive integer will be returned, else a negative integer will be returned.

## bwtekGatedMode

int bwtekGatedMode

(

int nGatedTime, 	// Gated Time between an External trigger and a CCD trigger

int nChannel  		// channel to get data from

);

This function is for controlling CCD gated delay time for specific models:

See Appendix H for list of spectrometer models which can use this function.

nGatedTime is used to determine the time delay before the CCD starts to acquire data after an external pulse has been delivered.  The parameter can be assigned an integer value of 0 to 65535.  This value is in addition to the inherent delay time.

Example:  When nGatedTime is 44 the gated time will be:

(9.2µs + 0.25µs*44) = 20.2µs  = (9.2µs + 0.25µs*nGatedTime)

**NOTE** When 0 is assigned to nGatedTime, the function will be disabled.

nChannel is used to address a specific spectrometer device to be operated when multiple spectrometer devices are involved.  Users must call the bwtekSetupChannel function to determine which channel numbers are available.  A total of 32 spectrometer devices may be connected at one time, where the nChannel value will range from 0 – 31.

RETURN

If the function call is successful, a positive integer will be returned, else a negative integer will be returned.

## bwtekSetExtPulse

int bwtekSetExtPulse

(

int nOnOff,			// switch to turn on or off pulse train

int nDelayTime,		// delay time between the CCD trigger and the first pulse

int nHigh, 			// time duration for high pulse width

int nLow, 			// time duration for low pulse width

int nPulse, 			// number of pulses to be generated

int nInverse,		// pulse output will be inverted

int nChannel  		// channel to get data from

);

To define output pulse, use this function first before acquiring data.

This function is used to generate a pulse train defined by nHigh, nLow and nPulse.
This bwteSetExtPulse function would be called Before the bwtekDataReadUSB function.

### Internal Trigger Mode

| USB Models | T_S (Software access time, const value) | T_Delay | T_High | T_Low |
| --- | --- | --- | --- | --- |
| BRC111A | 9.6us | 0.25us*(nDelayTime+1) | 0.25us*nHigh | 0.25us*nLow |
| BRC112P | 9.6us | 0.25us*(nDelayTime+1) | 0.25us*nHigh | 0.25us*nLow |
| BRC112E | 9.6us | 0.25us*(nDelayTime+1) | 0.25us*nHigh | 0.25us*nLow |
| BRC115E | N/A | N/A | N/A | N/A |
| BRC115U | N/A | N/A | N/A | N/A |
| BRC115V | N/A | N/A | N/A | N/A |
| BRC642E-2048 | 9.6us | 0.25us*(nDelayTime+1) | 0.25us*nHigh | 0.25us*nLow |
| BRC711E-512 | 9.6us | 0.25us*(nDelayTime+1) | 0.25us*nHigh | 0.25us*nLow |
| BRC711E-1024 | 9.6us | 0.25us*(nDelayTime+1) | 0.25us*nHigh | 0.25us*nLow |
| BRC741E-512 | 9.6us | 0.25us*(nDelayTime+1) | 0.25us*nHigh | 0.25us*nLow |
| BRC741E-1024 | 9.6us | 0.25us*(nDelayTime+1) | 0.25us*nHigh | 0.25us*nLow |
| BTC111E | 9.6us | 0.25us*(nDelayTime+1) | 0.25us*nHigh | 0.25us*nLow |
| BTC112E | 9.6us | 0.25us*(nDelayTime+1) | 0.25us*nHigh | 0.25us*nLow |
| BTC162E | 9.6us | 0.25us*(nDelayTime+1) | 0.25us*nHigh | 0.25us*nLow |
| BTC261E-256 | N/A | N/A | N/A | N/A |
| BTC261E-512 | N/A | N/A | N/A | N/A |
| BTC261E-1024 | N/A | N/A | N/A | N/A |
| BTC261P-512 | 9.6us | 0.25us*(nDelayTime+1) | 0.25us*nHigh | 0.25us*nLow |
| BTC262A-256 | 9.6us | 0.25us*(nDelayTime+1) | 0.25us*nHigh | 0.25us*nLow |
| BTC262E-256 | N/A | N/A | N/A | N/A |
| BTC262E-512 | N/A | N/A | N/A | N/A |
| BTC262P-512 | 9.6us | 0.25us*(nDelayTime+1) | 0.25us*nHigh | 0.25us*nLow |
| BTC263E-256 | 9.6us | 0.25us*(nDelayTime+1) | 0.25us*nHigh | 0.25us*nLow |
| BTC264P-256 | 9.6us | 0.25us*(nDelayTime+1) | 0.25us*nHigh | 0.25us*nLow |
| BTC264P-512 | 9.6us | 0.25us*(nDelayTime+1) | 0.25us*nHigh | 0.25us*nLow |
| BTC264P-1024 | 9.6us | 0.25us*(nDelayTime+1) | 0.25us*nHigh | 0.25us*nLow |
| BTC611E-512 | 9.6us | 0.25us*(nDelayTime+1) | 0.25us*(65536-nHigh) | 0.25us*(65536-nLow) |
| BTC611E-1024 | 9.6us | 0.25us*(nDelayTime+1) | 0.25us*(65536-nHigh) | 0.25us*(65536-nLow) |
| BTC613E-512 | 9.6us | 0.25us*(nDelayTime+1) | 0.25us*(65536-nHigh) | 0.25us*(65536-nLow) |
| BTC613E-1024 | 9.6us | 0.25us*(nDelayTime+1) | 0.25us*(65536-nHigh) | 0.25us*(65536-nLow) |
| BTC621E-512 | 9.6us | 0.25us*(nDelayTime+1) | 0.25us*(65536-nHigh) | 0.25us*(65536-nLow) |
| BTC621E-1024 | 9.6us | 0.25us*(nDelayTime+1) | 0.25us*(65536-nHigh) | 0.25us*(65536-nLow) |
| BTC651E-512 | 9.6us | 0.25us*(nDelayTime+1) | 0.25us*(65536-nHigh) | 0.25us*(65536-nLow) |
| BTC651E-1024 | 9.6us | 0.25us*(nDelayTime+1) | 0.25us*(65536-nHigh) | 0.25us*(65536-nLow) |
| BTC655E | N/A | N/A | N/A | N/A |
| BTC661E | 9.6us | 0.25us*(nDelayTime+1) | 0.25us*(65536-nHigh) | 0.25us*(65536-nLow) |
| BTC665E | N/A | N/A | N/A | N/A |

### External Trigger Mode

| USB Models | T_S (Software access time, it is const value) | T_Gate | T_Delay | T_High | T_Low |
| --- | --- | --- | --- | --- | --- |
| BRC111A | 9.6us | 0.25us*nGateTime | 0.25us*(nDelayTime+1) | 0.25us*nHigh | 0.25us*nLow |
| BRC112P | 9.6us | 0.25us*nGateTime | 0.25us*(nDelayTime+1) | 0.25us*nHigh | 0.25us*nLow |
| BRC112E | 9.6us | 0.25us*nGateTime | 0.25us*(nDelayTime+1) | 0.25us*nHigh | 0.25us*nLow |
| BRC115E | N/A | N/A | N/A | N/A | N/A |
| BRC115U | N/A | N/A | N/A | N/A | N/A |
| BRC115V | N/A | N/A | N/A | N/A | N/A |
| BRC642E-2048 | 9.6us | 0.25us*nGateTime | 0.25us*(nDelayTime+1) | 0.25us*nHigh | 0.25us*nLow |
| BRC711E-512 | 9.6us | 0.25us*nGateTime | 0.25us*(nDelayTime+1) | 0.25us*nHigh | 0.25us*nLow |
| BRC711E-1024 | 9.6us | 0.25us*nGateTime | 0.25us*(nDelayTime+1) | 0.25us*nHigh | 0.25us*nLow |
| BRC741E-512 | 9.6us | 0.25us*nGateTime | 0.25us*(nDelayTime+1) | 0.25us*nHigh | 0.25us*nLow |
| BRC741E-1024 | 9.6us | 0.25us*nGateTime | 0.25us*(nDelayTime+1) | 0.25us*nHigh | 0.25us*nLow |
| BTC111E | 9.6us | 0.25us*nGateTime | 0.25us*(nDelayTime+1) | 0.25us*nHigh | 0.25us*nLow |
| BTC112E | 9.6us | 0.25us*nGateTime | 0.25us*(nDelayTime+1) | 0.25us*nHigh | 0.25us*nLow |
| BTC162E | 9.6us | 0.25us*nGateTime | 0.25us*(nDelayTime+1) | 0.25us*nHigh | 0.25us*nLow |
| BTC261E-256 | N/A | N/A | N/A | N/A | N/A |
| BTC261E-512 | N/A | N/A | N/A | N/A | N/A |
| BTC261E-1024 | N/A | N/A | N/A | N/A | N/A |
| BTC261P-512 | 9.6us | 0.25us*nGateTime | 0.25us*(nDelayTime+1) | 0.25us*nHigh | 0.25us*nLow |
| BTC262A-256 | 9.6us | 0.25us*nGateTime | 0.25us*(nDelayTime+1) | 0.25us*nHigh | 0.25us*nLow |
| BTC262E-256 | N/A | N/A | N/A | N/A | N/A |
| BTC262E-512 | N/A | N/A | N/A | N/A | N/A |
| BTC262P-512 | 9.6us | 0.25us*nGateTime | 0.25us*(nDelayTime+1) | 0.25us*nHigh | 0.25us*nLow |
| BTC263E-256 | 9.6us | 0.25us*nGateTime | 0.25us*(nDelayTime+1) | 0.25us*nHigh | 0.25us*nLow |
| BTC264P-256 | 9.6us | 0.25us*nGateTime | 0.25us*(nDelayTime+1) | 0.25us*nHigh | 0.25us*nLow |
| BTC264P-512 | 9.6us | 0.25us*nGateTime | 0.25us*(nDelayTime+1) | 0.25us*nHigh | 0.25us*nLow |
| BTC264P-1024 | 9.6us | 0.25us*nGateTime | 0.25us*(nDelayTime+1) | 0.25us*nHigh | 0.25us*nLow |
| BTC611E-512 | 9.6us | 0.25us*nGateTime | 0.25us*(nDelayTime+1) | 0.25us*(65536-nHigh) | 0.25us*(65536-nLow) |
| BTC611E-1024 | 9.6us | 0.25us*nGateTime | 0.25us*(nDelayTime+1) | 0.25us*(65536-nHigh) | 0.25us*(65536-nLow) |
| BTC613E-512 | 9.6us | 0.25us*nGateTime | 0.25us*(nDelayTime+1) | 0.25us*(65536-nHigh) | 0.25us*(65536-nLow) |
| BTC613E-1024 | 9.6us | 0.25us*nGateTime | 0.25us*(nDelayTime+1) | 0.25us*(65536-nHigh) | 0.25us*(65536-nLow) |
| BTC621E-512 | 9.6us | 0.25us*nGateTime | 0.25us*(nDelayTime+1) | 0.25us*(65536-nHigh) | 0.25us*(65536-nLow) |
| BTC621E-1024 | 9.6us | 0.25us*nGateTime | 0.25us*(nDelayTime+1) | 0.25us*(65536-nHigh) | 0.25us*(65536-nLow) |
| BTC651E-512 | 9.6us | 0.25us*nGateTime | 0.25us*(nDelayTime+1) | 0.25us*(65536-nHigh) | 0.25us*(65536-nLow) |
| BTC651E-1024 | 9.6us | 0.25us*nGateTime | 0.25us*(nDelayTime+1) | 0.25us*(65536-nHigh) | 0.25us*(65536-nLow) |
| BTC655E | N/A | N/A | N/A | N/A | N/A |
| BTC661E | 9.6us | 0.25us*nGateTime | 0.25us*(nDelayTime+1) | 0.25us*(65536-nHigh) | 0.25us*(65536-nLow) |
| BTC665E | N/A | N/A | N/A | N/A | N/A |

nOnOff will turn on or off the output pulse train.  If 0 is assigned, the output signal train will not be generated (no signal).  If 1 is assigned, the output signal train will be synchronized with the CCD trigger pulse that is generated after the normal delay time of (0.25us*(nDelayTime+1)).

nDelayTime is used to determine the delay time between the CCD trigger and the first pulse.  This parameter can be assigned an integer value of 1 to 65535.  This value is in addition to the inherent delay time.  Example:  When nDelayTime is 50, the total delay time will be:

12.75us=(0.25us*(nDelayTime+1))

nHigh is used to determine the high pulse width. The parameter can be assigned an integer value of 1 to 65535.  This value is in addition to the inherent delay time.
Example:  
BTC611E - When nHigh is 63, the high pulse width will be 16368.25us=0.25us*(65536-nHigh)

Other (BTC112, BTC261, BTC641 etc.) - When nHigh is 63, the high pulse width will be 15.75us=(0.25us*nHigh)

nLow is used to determine the low pulse width. The parameter can be assigned an integer value of 1 to 65535.  This value is in addition to the inherent delay time.

Example:  
BTC611E - When nLow is 63, the low pulse width will be 16368.25us=0.25us*(65536-nLow)

Other (BTC112, BTC261, BTC641 etc.) - When nLow is 63, the low pulse width will be 15.75us=(0.25us*nLow)

nPulse is used to set the number of pulses that are to be generated.  The parameter can be assigned an integer value of 1 to 65535.

**NOTE**  In the function bwtekSetTime the value assigned to the lTime parameter must be longer

than (T_S+T_Gate+T_Delay+(T_High+T_Low)*nPulse) if using external trigger.  Or must be longer

than (T_S+T_Delay+(T_High+T_Low)*nPulse) if using internal trigger.

nInverse is used to invert the output pulse train.  1 will output the inversed pulse train.  0 will output normal pulse train.

nChannel is used to address a specific spectrometer device to be operated when multiple spectrometer devices are involved.  Users must call the bwtekSetupChannel function to determine which channel numbers are available.  A total of 32 spectrometer devices may be connected at one time, where the nChannel value will range from 0 – 31.

RETURN

If the function call is successful, a positive integer will be returned, else a negative integer will be returned.

# AUX Port Functions 2---Shutter Functions:

For older spectrometers, in their user manual they may have different names for TTL Input, such as TTL Sensing.   The term TTL Sensing and TTL Input are one in the same.

**Reference the Appendices, SDK Function Groups, TTL Output and TTL Input to verify if your spectrometer model can use the below function calls.

The following shutter functions work for specific model types which have shutters installed.

## bwtekShutterOpen (OBSOLETE)

int bwtekShutterOpen

(

int nChannel  	  // channel to get data from

);

Please use the below function, bwtekShutterControl in place of this function.

## bwtekShutterClose (OBSOLETE)

int bwtekShutterClose

(

int nChannel  	  // channel to get data from

);

Please use the below function, bwtekShutterControl in place of this function.

## bwtekShutterControl

int bwtekShutterControl

(

int nSetShutter1,	// sets Shutter 1 state

int nSetShutter2,	// sets Shutter 2 state

int nChannel  		// channel to get data from

);

This function controls the ‘Shutter 1’ and ‘Shutter 2’ of a Spectrometer’s AUX Port pin assignment.
*Note: Check your spectrometer’s User Manual Aux Pin assignment to see if ‘Shutter 1’ is a valid Pin Assignment.  If ‘Shutter 1’ does not appear, this function will Not work with your spectrometer.

nSetShutter1 is used to set Shutter 1 to open or close.

0 = Close Shutter, 1 = Open Shutter, -1 = No Action.

nSetShutter2 is used to set Shutter 2 to open or close.

0 = Close Shutter, 1= Open Shutter, -1 = No Action.

nChannel is used to address a specific spectrometer device to be operated when multiple spectrometer devices are involved.  Users must call the bwtekSetupChannel function to determine which channel numbers are available.  A total of 32 spectrometer devices may be connected at one time, where the nChannel value will range from 0 – 31.

RETURN

If the function call is successful, a positive integer will be returned, else a negative integer will be returned.

# AUX Port Functions 3

For older spectrometers, in their user manual they may have different names for TTL Input, such as TTL Sensing.   The term TTL Sensing and TTL Input are one in the same.

**Reference the Appendices, SDK Function Groups, TTL Output and TTL Input to verify if your spectrometer model can use the below function calls.

## bwtekGetExtStatus

int bwtekGetExtStatus

(

int nChannel  	//channel to get data from

);

This function is used to retrieve the status of the TTL Input pin.

nChannel is used to address a specific spectrometer device to be operated when multiple spectrometer devices are involved.  Users must call the bwtekSetupChannel function to determine which channel numbers are available.  A total of 32 spectrometer devices may be connected at one time, where the nChannel value will range from 0 – 31.

RETURN

If the status of input pin is low, a 0 will be returned. If the status of input pin is high, a positive integer will be returned.  If the function call is unsuccessful, a negative integer will be returned.

## bwtekGetTTLIn

int bwtekGetTTLIn

(

int nNo,      		// The ‘TTL Input Number’ to use from the spectrometer

int *nTTLStatus,   	// Return value of TTL Input Signal

int nChannel  	    	// Channel

);

This function is used to retrieve the status of the selected TTL input pin.

(Reference the Spectrometer’s User Manual for Pin location)

nNo is used to select which TTL Input Number to select from the spectrometer’s AUX port to use.  The default value should be 0 when the spectrometer has only a single TTL Input .  If there are more than one TTL Input Pin then for TTL Input 1/ Digital Input 1, nNo = 0.  If using the second TTL Input a value of 1 should be used.
*NOTE: nNo is Not related to the actual Pin Number.

(Reference the Spectrometer’s User Manual for Pin definition)

*nTTLStatus is a pointer to an array which will return the signal status of the selected TTL Input, where 1 = High and 0 = Low.

nChannel is used to address a specific spectrometer device to be operated when multiple spectrometer devices are involved.  Users must call the bwtekSetupChannel function to determine which channel numbers are available.  A total of 32 spectrometer devices may be connected at one time, where the nChannel value will range from 0 – 31.

RETURN

If the function call is successful, a positive integer will be returned, else a negative integer will be returned.

## bwtekSetTTLOut

int bwtekSetTTLOut

(

int nNo,       		// The ‘TTL Output Number’ to use from the spectrometer

int nTTLStatus,		// Set the TTL output status

int nInverse,  		// TTL pin signal inversed

int nChannel  		// Channel

);

This function is used to set the status of the selected TTL Output pin to: On / Off (High / Low).

(Reference the Spectrometer’s User Manual for Pin location/definition)

nNo is used to select which TTL Output Number from the spectrometer’s AUX port to use.  The default value should be 0 when the spectrometer has only a single TTL Output.  If there are more than one TTL Output Pin then for TTL Output 1 / Digital Output 1 nNo = 0.  If using the second TTL Output a value of 1 should be used.
*NOTE: nNo is Not related to the actual Pin Number.

(Reference the Spectrometer’s User Manual for Pin definition)

nTTLStatus is used to set the status of an TTL output signal.  It can be set from 0 to 1. (0 = Low, 1 = High)

nInverse is used to invert the output pin signal for nTTLStatus.  1 will invert the nTTLStatus values, and 0 will keep them at default.

nChannel is used to address a specific spectrometer device to be operated when multiple spectrometer devices are involved.  Users must call the bwtekSetupChannel function to determine which channel numbers are available.  A total of 32 spectrometer devices may be connected at one time, where the nChannel value will range from 0 – 31.

RETURN

If the function call is successful, a positive integer will be returned, else a negative integer will be returned.

## bwtekGetAnalogIn

int bwtekGetAnalogIn

(

int nNo,     		// Analog Input Number

int *nADValue,		// AD value

double *lValtage, 	// Voltage value

int nChannel  		// Channel

);

This function is used to get the value of analog signal input (Analog Sensing).

nNo is used to select the Analog Input Number feature to select from the spectrometer’s AUX port.  The default value should be 0 when the spectrometer has only a single TTL Output or selecting the first TTL Output pin.  If using the second TTL Output feature a value of 1 should be used.
*NOTE this Value is Not related to the actual Pin Number.

nADValue is A/D Value of analog input. Valid range is 0...4095 (12 bit)

lVoltage is voltage of analog input. Valid range is 0…2.5 (V)

nChannel is used to address a specific spectrometer device to be operated when multiple spectrometer devices are involved.  Users must call the bwtekSetupChannel function to determine which channel numbers are available.  A total of 32 spectrometer devices may be connected at one time, where the nChannel value will range from 0 – 31.

RETURN

If the function call is successful, a positive integer will be returned, else a negative integer will be returned.

# BTC261E & BTC262E Functions:

The following functions work only with the BTC261E & BTC262E spectrometers.

## bwtekSetABGain (Obsolete)

* Obsolete Note* Units built before 2008 can still use this function.

Units built around 2008 and later this function will not work.

The ability to control the A and B Channels of the detector was changed from Software Control to Hardware Control.  If your spectrometer has at least two trimpots on the front panel, at the bottom left corner from the SMA Plate then your unit has Hardware Control.  If you are still unsure if you can use this function, please contact B&W Tek, Inc. (www.bwtek.com)

int bwtekSetABGain

(

int nAB,		// select between the odd(A) and even(B) pixel channels

int nGain,		// set the gain value

int nChannel 	// channel to get data from

);

This function is for setting the gain value of the analog output on either A or B channel.

nAB is used to choose between the odd(A) and even(B) pixel channels.  Use 0 to control the odd(A) pixel channel and use 1 to control the even(B) pixel channel.

**NOTE For the BTC262E nAB must always be 0.  The BTC262E uses one channel for both odd and even pixels.

nGain is used to set the gain value of an analog output signal.  It can be set from 1 to 1023.

nChannel is used to address a specific spectrometer device to be operated when multiple spectrometer devices are involved.  Users must call the bwtekSetupChannel function to determine which channel numbers are available.  A total of 32 spectrometer devices may be connected at one time, where the nChannel value will range from 0 – 31.

RETURN

If the function call is successful, a positive integer will be returned, else a negative interger will be returned.

Notes: When this function is used, the difference between desired set value and current value should less than 90. For example, the current value that you read out is 100, you want to set 1000. So you should set it as following code:

Dim OldValue as integer

Dim NewValue as integer

Dim tmp_Count as integer

Dim i ,channel as integer

Channel=0

OldValue=100

NewValue=1000

tmp_count=Int (abs(NewValue-OldValue+1))/90

For i=1 to tmp_count

tmp_ret:=bwtekSetABGain(0,OldValue+ i*90,channel)

Next i

tmp_ret=bwtekSetABGain(0, NewValue, channel)

## bwtekSetABOffset (Obsolete)

* Obsolete Note* Units built before 2008 can still use this function.

Units built around 2008 and later this function will not work.

The ability to control the A and B Channels of the detector was changed from Software Control to Hardware Control.  If your spectrometer has at least two trimpots on the front panel, at the bottom left corner from the SMA Plate then your unit has Hardware Control.  If you are still unsure if you can use this function, please contact B&W Tek, Inc. (www.bwtek.com)

int bwtekSetABOffset

(

int nAB,		// select between the odd(A) and the even(B) pixel channels

int nOffset,		// set the offset value

int nChannel  	// channel to get data from

);

This function is for setting the offset value of the analog output on either A or B channel

nAB is used to choose between the odd(A) and even(B) pixel channels.  Use 0 to control the odd (A) pixel channel and use 1 to control the even (B) pixel channel.

**NOTE For the BTC262E nAB must always be 0.  The BTC262E uses one channel for both odd and even pixels.

nOffset is used to set the offset value of an analog output signal.  It can be set from 1 to 1023.

nChannel is used to address a specific spectrometer device to be operated when multiple spectrometer devices are involved.  Users must call the bwtekSetupChannel function to determine which channel numbers are available.  A total of 32 spectrometer devices may be connected at one time, where the nChannel value will range from 0 – 31.

RETURN

If the function call is successful, a positive integer will be returned, else a negative integer will be returned.

Notes: when calling this function, the difference between desired set value and current value should less than 90. For example, if the current readout value is 100 and you want to set it to 1000, use the following code:

Dim OldValue as integer

Dim NewValue as integer

Dim tmp_Count as integer

Dim i ,channel as integer

Channel=0

OldValue=100

NewValue=1000

tmp_count=Int (abs(NewValue-OldValue+1))/90

For i=1 to tmp_count

tmp_ret:= bwtekSetABOffset (0,OldValue+ i*90,channel)

Next i

tmp_ret= bwtekSetABOffset (0, NewValue, channel)

## bwtekGetABGain (Obsolete)

* Obsolete Note* Units built before 2008 can still use this function.

Units built around 2008 and later this function will not work.

The ability to control the A and B Channels of the detector was changed from Software Control to Hardware Control.  If your spectrometer has at least two trimpots on the front panel, at the bottom left corner from the SMA Plate then your unit has Hardware Control.  If you are still unsure if you can use this function, please contact B&W Tek, Inc. (www.bwtek.com)

int bwtekGetABGain

(

int nAB,		// select between the odd(A) and even(B) pixel channels

int *nGain,		// return value of gain

int nChannel 	// channel to get data from

);

This function is for getting the gain value of the analog output on either A or B channel

nAB is used to choose between the odd(A) and even(B) pixel channels. Use 0 to control the odd (A) pixel channel and use 1 to control the even (B) pixel channel.

**NOTE For the BTC262E nAB must always be 0. The BTC262E uses one channel for both odd and even pixels.

*nGain is a pointer to an array that will return the gain value of an analog output signal.

nChannel is used to address a specific spectrometer device to be operated when multiple spectrometer devices are involved.  Users must call the bwtekSetupChannel function to determine which channel numbers are available.  A total of 32 spectrometer devices may be connected at one time, where the nChannel value will range from 0 – 31.

RETURN

If the function call is successful, a positive integer will be returned, else a negative integer will be returned.

## bwtekGetABOffset (Obsolete)

* Obsolete Note* Units built before 2008 can still use this function.

Units built around 2008 and later this function will not work.

The ability to control the A and B Channels of the detector was changed from Software Control to Hardware Control.  If your spectrometer has at least two trimpots on the front panel, at the bottom left corner from the SMA Plate then your unit has Hardware Control.  If you are still unsure if you can use this function, please contact B&W Tek, Inc. (www.bwtek.com)

int bwtekGetABOffset

(

int nAB,		// select between the odd(A) and even(B) pixel channels

int *nOffset,		// return value of offset

int nChannel 	// channel to get data from

);

This function is for getting the offset value of the analog output on either A or B channel

nAB is used to choose between the odd( A) and even (B) pixel channels.  Use 0 to control the odd (A) pixel channel and use 1 to control the even (B) pixel channel.

**NOTE For the BTC262E nAB must always be 0.  The BTC262E uses one channel for both odd and even pixels.

*nOffset is a pointer to an array that will return the offset value of an analog output signal.

nChannel is used to address a specific spectrometer device to be operated when multiple spectrometer devices are involved.  Users must call the bwtekSetupChannel function to determine which channel numbers are available.  A total of 32 spectrometer devices may be connected at one time, where the nChannel value will range from 0 – 31.

RETURN

If the function call is successful, a positive integer will be returned, else a negative integer will be returned.

## bwtekSetInGaAsMode

int bwtekSetInGaAsMode

(

int nMode,		// select feedback capacitor mode

int nChannel  	// channel to get data from

);

This function is for choosing the detector mode: high dynamic range or high sensitivity range.

nMode is used to choose the feedback capacitor mode of the detector.

| nMode | Detector Mode |
| --- | --- |
| 0 | High Sensitivity |
| 1 | High Dynamic Range |

nChannel is used to address a specific spectrometer device to be operated when multiple spectrometer devices are involved.  Users must call the bwtekSetupChannel function to determine which channel numbers are available.  A total of 32 spectrometer devices may be connected at one time, where the nChannel value will range from 0 – 31.

RETURN

If the function call is successful, a positive integer will be returned, else a negative integer will be returned.

## bwtekGetInGaAsMode

int bwtekGetInGaAsMode

(

int *nMode,		// select detector mode

int nChannel  	// channel to get data from

);

This function is for retrieving the detector mode setting: high dynamic range or high sensitivity range.

*nMode is a pointer to an array that will return the value of the feedback capacitor mode setting.

| *nMode | Detector Mode |
| --- | --- |
| 0 | High Sensitivity |
| 1 | High Dynamic Range |

nChannel is used to address a specific spectrometer device to be operated when multiple spectrometer devices are involved.  Users must call the bwtekSetupChannel function to determine which channel numbers are available.  A total of 32 spectrometer devices may be connected at one time, where the nChannel value will range from 0 – 31.

RETURN

If the function call is successful, a positive integer will be returned, else a negative integer will be returned.

## bwtekQueryTemperature

int bwtekQueryTemperature

(

int nCommand,

int *nReserved,

double *nTempValue,

int nChannel

);

This function is for reading the temperature of the CCD Detector OR temperature of the sample holder.

**Note your spectrometer hardware must be compatible for this function to work.

If this function is not working for you correct, contact BWTEK to verify that your spectrometer has compatible hardware for this function.

nCommand is a flag of specific temperature. 11 for BTC261E CCD detector temperature, 81 for BTC261E sample holder temperature.

*nReserved is a pointer to an integer return value. It is reserved.

*nTempValue is a pointer to float pointer value of temperature. The temperature units are degrees Celsius.

nChannel is used to address a specific spectrometer device to be operated when multiple spectrometer devices are involved.  Users must call the bwtekSetupChannel function to determine which channel numbers are available.  A total of 32 spectrometer devices may be connected at one time, where the nChannel value will range from 0 – 31.

RETURN

If the function call is successful, a positive integer will be returned, else a negative integer will be returned.

## bwtekAccessDeltaTemp

int bwtekAccessDeltaTemp

(

int nReadWrite,

double *pDeltaTempValue,

int nChannel

);

This function is for read/write the adjust value of temperature of sample holder and CCD Detector.

nReadWrite is a flag of Read or write. 0 for read delta of temperature, 1 for write delta of temperature.

*nDeltaTempValue is a pointer to float pointer value of temperature delta. The temperature units are degrees Celsius. The real temperature is sum of nDeltaTempValue and nTempValue

nChannel is used to address a specific spectrometer device to be operated when multiple spectrometer devices are involved.  Users must call the bwtekSetupChannel function to determine which channel numbers are available.  A total of 32 spectrometer devices may be connected at one time, where the nChannel value will range from 0 – 31.

RETURN

If the function call is successful, a positive integer will be returned, else a negative integer will be returned.

## bwtekAccessDeltaTemp1

int bwtekAccessDeltaTemp1

(

int nReadWrite,

double *pDeltaTempValue,

double *pDeltaTempValue1,

int nChannel

);

This function is for read/write the adjust value of temperature of sample holder and CCD Detector.

nReadWrite is a flag of Read or write. 0 for read delta of temperature, 1 for write delta of temperature.

*nDeltaTempValue is a pointer to float pointer value of temperature delta of sample holder. The temperature units are degrees Celsius. The real temperature is sum of nDeltaTempValue and nTempValue.  Please see notes for detail information.

*nDeltaTempValue1 is a pointer to float pointer value of temperature delta of BTC261E/BTC262E CCD detector. The temperature units are degrees Celsius. The real temperature is sum of nDeltaTempValue1 and nTempValue1 . Please see notes for detail information.

nChannel is used to address a specific spectrometer device to be operated when multiple spectrometer devices are involved.  Users must call the bwtekSetupChannel function to determine which channel numbers are available.  A total of 32 spectrometer devices may be connected at one time, where the nChannel value will range from 0 – 31.

RETURN

If the function call is successful, a positive integer will be returned, else a negative integer will be returned.

Notes:

- For BTC261E/BTC262E sample holder temperature

Call function bwtekQueryTemperature (81, nReserved, nTempValue, channel), get value nTempValue

Call function BwtekAccessDeltaTemp1 (0, nDeltaTempValue, nDeltaTempValue1, channel), get value nDeltaTempValue

The BTC261E/BTC262E sample holder temperature equal (nTempValue + nDeltaTempValue)

- For BTC261E/BTC262E CCD detector temperature

Call function bwtekQueryTemperature (11, nReserved, nTempValue1, channel), get value nTempValue1

Call function BwtekAccessDeltaTemp1 (0, nDeltaTempValue, nDeltaTempValue1, channel), get value nDeltaTempValue1

The BTC261E/BTC262E CCD detector temperature equal (nTempValue1 + nDeltaTempValue1)

## bwtekWriteValue

int bwtekWriteValue

(

int nItem,

int nSetValue,

int nChannel

);

This function is to set the parameter of integration time.

nItem control which variable for integration time setting to change.

0 is for setting the Integration time units.

1 is for setting the integration time multiplier.

2 is for setting the integration time.

nSetValue is a set value for the ‘nItem’ parameter used.

When nItem = 0, use an ‘nSetValue’ of 0 for microseconds and 1 is for milliseconds.

nChannel is used to address a specific spectrometer device to be operated when multiple spectrometer devices are involved.  Users must call the bwtekSetupChannel function to determine which channel numbers are available.  A total of 32 spectrometer devices may be connected at one time, where the nChannel value will range from 0 – 31.

For example:

Value0=0

bwtekWriteValue(0, Value0, channel);  //Set the unit of integration time as microseconds (us).

Value0=1

bwtekWriteValue(0, Value0, channel);  //Set the unit of integration time as milliseconds (ms).

Value0=2

bwtekWriteValue(1, Value0, channel);  //Set the multiply of integration time as 2

Value0=34

bwtekWriteValue(2, Value0, channel);  //Set the integration time to 34.

RETURN

If the function call is successful, a positive integer will be returned, else a negative integer will be returned.

## bwtekReadValue

int bwtekReadValue

(

int nItem,

int *nGetValue,

int nChannel

);

This function is for read/write the adjust value of temperature of sample holder and CCD Detector.

nItem control which variable for integration time setting to read.

0 is for reading the Integration time units.

1 is for reading the integration time multiplier.

2 is for reading the integration time.

*nGetValue is a pointer to return value of desired parameter of integration time.

nChannel is used to address a specific spectrometer device to be operated when multiple spectrometer devices are involved.  Users must call the bwtekSetupChannel function to determine which channel numbers are available.  A total of 32 spectrometer devices may be connected at one time, where the nChannel value will range from 0 – 31.

For example:

bwtekReadValue (0, &Value0, channel); 	// Value0 is the integration time unit.  value0 = 0 = microseconds (us), // if value0 = 1 = milliseconds (ms)

bwtekReadValue (1, &Value0, channel);	// Value0 is the multiplier value.

bwtekReadValue (2, &Value0, channel); 	// the Value0 is the integration value.

RETURN

If the function call is successful, a positive integer will be returned, else a negative integer will be returned.

# BTC262A &BTC263E Functions:

The following functions work with the BTC263E and BTC262A spectrometers:

## bwtekSetTimeUnitUSB

int bwtekSetTimeUnitUSB

(

int nTimeUnit,

int nChannel

);

This function is for setting the integration time ‘unit’.

nTimeUnit is a unit of integration time, 0 for microsecond (us), 1 for millisecond (ms).

nChannel is used to address a specific spectrometer device to be operated when multiple spectrometer devices are involved.  Users must call the bwtekSetupChannel function to determine which channel numbers are available.  A total of 32 spectrometer devices may be connected at one time, where the nChannel value will range from 0 – 31.

RETURN

Is the function call is successful the time nTimeUnit value will be returned else a negative integer will be returned.

## bwtekGetTimeUnitUSB

int bwtekGetTimeUnitUSB

(

int *TimeUnit,

int nChannel

);

This function is for reading the integration time ‘unit.’

*nTimeUnit is a pointer to an array that will return the unit value for the integration time, 0 for microsecond (us), 1 for millisecond (ms).

nChannel is used to address a specific spectrometer device to be operated when multiple spectrometer devices are involved.  Users must call the bwtekSetupChannel function to determine which channel numbers are available.  A total of 32 spectrometer devices may be connected at one time, where the nChannel value will range from 0 – 31.

RETURN

Is the function call is successful the time the time ‘unit’ will be returned else a negative integer will be returned.

## bwtekSetInGaAsMode

int bwtekSetInGaAsMode

(

int nSetModeValue,		// select detector mode

int nChannel  			// channel to get data from

);

This function is for setting the detector mode.

nSetModeValue is used to choose detector mode .

| nSetModeValue | Detector Mode |
| --- | --- |
| 0x00 | Maximum Sensitivity (default) |
| 0x80 | High Sensitivity |
| 0x180 | High Dynamic Range |
| 0x100 | Maximum Dynamic Range |

nChannel is used to address a specific spectrometer device to be operated when multiple spectrometer devices are involved.  Users must call the bwtekSetupChannel function to determine which channel numbers are available.  A total of 32 spectrometer devices may be connected at one time, where the nChannel value will range from 0 – 31.

RETURN

If the function call is successful, a positive integer will be returned, else a negative integer will be returned.

## bwtekGetInGaAsMode

int bwtekGetInGaAsMode

(

int *nMode,		// select detector mode

int nChannel  	// channel to get data from

);

This function is for retrieving the detector mode setting.

*Prerequisite Note: The function call ‘bwtekSetInGaAsMode’ must be called first for this function to work.

*nMode is a pointer to an array that will return the value of the feedback capacitor mode setting

| nSetModeValue | Detector Mode |
| --- | --- |
| 0x00 | Maximum Sensitivity (default) |
| 0x80 | High Sensitivity |
| 0x180 | High Dynamic Range |
| 0x100 | Maximum Dynamic Range |

nChannel is used to address a specific spectrometer device to be operated when multiple spectrometer devices are involved.  Users must call the bwtekSetupChannel function to determine which channel numbers are available.  A total of 32 spectrometer devices may be connected at one time, where the nChannel value will range from 0 – 31.

RETURN

If the function call is successful, a positive integer will be returned, else a negative integer will be returned.

# BTC261P & BTC262P & BTC264P Functions:

The following functions work with the BTC261P & BTC262P & BTC264P spectrometers:

## bwtekSetTimeUnitUSB (OBSOLETE)

int bwtekSetTimeUnitUSB

(

int nTimeUnit,

int nChannel

);

This function is for setting the integration time ‘unit’.

*OBOLSETE NOTE*

The integration time unit range for these model spectrometers are microseconds (us).

Range: 200us – 2,100,000,000us

## bwtekGetTimeUnitUSB (OBSOLETE)

int bwtekGetTimeUnitUSB

(

int *TimeUnit,

int nChannel

);

This function is for reading the integration time ‘unit.’

*OBOLSETE NOTE*

The integration time unit for these model spectrometers will always be microseconds (us).

Range: 200us – 2,100,000,000us

## bwtekSetInGaAsMode

int bwtekSetInGaAsMode

(

int nSetModeValue,		// select detector mode

int nChannel  			// channel to get data from

);

This function is for setting the detector mode.

nSetModeValue is used to choose detector mode .

| nSetModeValue | Detector Mode |
| --- | --- |
| 0 | High Sensitivity |
| 1 | High Dynamic Range |

nChannel is used to address a specific spectrometer device to be operated when multiple spectrometer devices are involved.  Users must call the bwtekSetupChannel function to determine which channel numbers are available.  A total of 32 spectrometer devices may be connected at one time, where the nChannel value will range from 0 – 31.

RETURN

If the function call is successful, a positive integer will be returned, else a negative integer will be returned.

## bwtekGetInGaAsMode

int bwtekGetInGaAsMode

(

int *nMode,		// select detector mode

int nChannel  	// channel to get data from

);

This function is for retrieving the detector mode setting.

*Prerequisite Note: The function call ‘bwtekSetInGaAsMode’ must be called first for this function to work.

*nMode is a pointer to an array that will return the value of the feedback capacitor mode setting

| *nSetModeValue | Detector Mode |
| --- | --- |
| 0 | High Sensitivity |
| 1 | High Dynamic Range |

nChannel is used to address a specific spectrometer device to be operated when multiple spectrometer devices are involved.  Users must call the bwtekSetupChannel function to determine which channel numbers are available.  A total of 32 spectrometer devices may be connected at one time, where the nChannel value will range from 0 – 31.

RETURN

If the function call is successful, a positive integer will be returned, else a negative integer will be returned.

# Appendix A: Model Descriptions

| USB Model | Description | Digitizer Resolution (bits) |
| --- | --- | --- |
| BRC111A | Non-Cooled 2048 CCD array spectrometer | 16 |
| BRC112P | Non-Cooled 2048 CCD array spectrometer | 16 |
| BRC112E | Non-Cooled 2048 CCD array spectrometer | 16 |
| BRC115E | Non-Cooled 2048 CCD array spectrometer | 16 |
| BRC115U | Non-Cooled 2048 CCD array spectrometer | 16 |
| BRC115V | Non-Cooled 2048 CCD array spectrometer | 16 |
| BRC642E-2048 | Non-Cooled 2048 Back Thinned CCD | 16 |
| BRC711E-512 | Non-Cooled 512 PDA array spectrometer | 16 |
| BRC711E-1024 | Non-Cooled 1024 PDA array spectrometer | 16 |
| BRC741E-512 | Non-Cooled 512 PDA array spectrometer | 16 |
| BRC741E-1024 | Non-Cooled 1024 PDA array spectrometer | 16 |
| BTC111E | TE Cooled 2048 CCD array spectrometer | 16 |
| BTC112E | TE Cooled 2048 CCD array spectrometer | 16 |
| BTC162E | TE Cooled 2048 CCD array spectrometer | 16 |
| BTC261E-256 | TE Cooled 256 InGaAs array spectrometer | 16 |
| BTC261E-512 | TE Cooled 512 InGaAs array spectrometer | 16 |
| BTC261E-1024 | TE Cooled 1024 InGaAs array spectrometer | 16 |
| BTC261P-512 | TE Cooled 512 InGaAs array spectrometer | 16 |
| BTC262A-256 | TE Cooled 256 InGaAs array spectrometer | 16 |
| BTC262E-256 | TE Cooled 256 InGaAs array spectrometer | 16 |
| BTC262E-512 | TE Cooled 512 InGaAs array spectrometer | 16 |
| BTC262P-512 | TE Cooled 512 InGaAs array spectrometer | 16 |
| BTC263E-256 | TE Cooled 256 InGaAs array spectrometer | 16 |
| BTC264P-256 | TE Cooled 512 InGaAs array spectrometer | 16 |
| BTC264P-512 | TE Cooled 512 InGaAs array spectrometer | 16 |
| BTC264P-1024 | TE Cooled 512 InGaAs array spectrometer | 16 |
| BTC611E-512 | TE Cooled 512 Back Thinned CCD | 16 |
| BTC611E-1024 | TE Cooled 1024 Back Thinned CCD | 16 |
| BTC613E-512 | TE Cooled 512 Back Thinned CCD | 16 |
| BTC613E-1024 | TE Cooled 1024 Back Thinned CCD | 16 |
| BTC621E-512 | TE Cooled 512 Back Thinned CCD | 16 |
| BTC621E-1024 | TE Cooled 1024 Back Thinned CCD | 16 |
| BTC651E-512 | TE Cooled 512 Back Thinned CCD | 16 |
| BTC651E-1024 | TE Cooled 1024 Back Thinned CCD | 16 |
| BTC655E | TE Cooled 2048 Back Thinned CCD | 16 |
| BTC661E | TE Cooled 2048 Back Thinned CCD | 16 |
| BTC665E | TE Cooled 2048 Back Thinned CCD | 16 |

# Appendix B: USB Interface Capabilities

| USB Models | USB 2.0 Connector: Uses FX2 USB Chip | USB 3.0 Connector: Uses FX3 USB Chip |
| --- | --- | --- |
| BRC111A | Yes |  |
| BRC112P | Yes |  |
| BRC112E | Yes |  |
| BRC115E |  | Yes |
| BRC115U |  | Yes |
| BRC115V |  | Yes |
| BRC642E-2048 | Yes |  |
| BRC711E-512 | Yes |  |
| BRC711E-1024 | Yes |  |
| BRC741E-512 | Yes |  |
| BRC741E-1024 | Yes |  |
| BTC111E | Yes |  |
| BTC112E | Yes |  |
| BTC162E | Yes |  |
| BTC261E-256 | Yes |  |
| BTC261E-512 | Yes |  |
| BTC261E-1024 | Yes |  |
| BTC261P-512 | Yes |  |
| BTC262A-256 | Yes |  |
| BTC262E-256 | Yes |  |
| BTC262E-512 | Yes |  |
| BTC262P-512 | Yes |  |
| BTC263E-256 | Yes |  |
| BTC264P-256 | Yes |  |
| BTC264P-512 | Yes |  |
| BTC264P-1024 | Yes |  |
| BTC611E-512 | Yes |  |
| BTC611E-1024 | Yes |  |
| BTC613E-512 | Yes |  |
| BTC613E-1024 | Yes |  |
| BTC621E-512 | Yes |  |
| BTC621E-1024 | Yes |  |
| BTC651E-512 | Yes |  |
| BTC651E-1024 | Yes |  |
| BTC655E |  | Yes |
| BTC661E | Yes |  |
| BTC665E |  | Yes |

# Appendix C: Spectrometer Driver Files

| USB Models | VID (Hex) | PID (Hex) | SPT File | INF File | DLL File |
| --- | --- | --- | --- | --- | --- |
| BRC111A | 16a3 | 2ec7 | bwtekusb.spt | bwtekusb2.inf | bwtekusb.dll |
| BRC112P | 16a3 | 2ec7 | bwtekusb.spt | bwtekusb2.inf | bwtekusb.dll |
| BRC112E | 16a3 | 2ec7 | bwtekusb.spt | bwtekusb2.inf | bwtekusb.dll |
| BRC115E | 16a3 | 2ed0 | BRC115.spt | bwtekusb3.inf | bwtekusb.dll |
| BRC115U | 16a3 | 2ed0 | BRC115.spt | bwtekusb3.inf | bwtekusb.dll |
| BRC115V | 16a3 | 2ed0 | BRC115.spt | bwtekusb3.inf | bwtekusb.dll |
| BRC642E-2048 | 16a3 | 2ecd | bwteku14.spt | bwtekusb2.inf | bwtekusb.dll |
| BRC711E-512 | 16a3 | 2ec7 | bwtekusb.spt | bwtekusb2.inf | bwtekusb.dll |
| BRC711E-1024 | 16a3 | 2ec7 | bwtekusb.spt | bwtekusb2.inf | bwtekusb.dll |
| BRC741E-512 | 16a3 | 2ec7 | bwtekusb.spt | bwtekusb2.inf | bwtekusb.dll |
| BRC741E-1024 | 16a3 | 2ec7 | bwtekusb.spt | bwtekusb2.inf | bwtekusb.dll |
| BTC111E | 16a3 | 2ec7 | bwtekusb.spt | bwtekusb2.inf | bwtekusb.dll |
| BTC112E | 16a3 | 2ec7 | bwtekusb.spt | bwtekusb2.inf | bwtekusb.dll |
| BTC162E | 16a3 | 2ec7 | bwtekusb.spt | bwtekusb2.inf | bwtekusb.dll |
| BTC261E-256 | 16a3 | 2ec4 | bwtekus6.spt | bwtekusb2.inf | bwtekusb.dll |
| BTC261E-512 | 16a3 | 2ec4 | bwtekus6.spt | bwtekusb2.inf | bwtekusb.dll |
| BTC261E-1024 | 16a3 | 2ec4 | bwtekus6.spt | bwtekusb2.inf | bwtekusb.dll |
| BTC261P-512 | 16a3 | 2ecf | bwteku16.spt | bwtekusb2.inf | bwtekusb.dll |
| BTC262A-256 | 16a3 | 2ece | bwteku15.spt | bwtekusb2.inf | bwtekusb.dll |
| BTC262E-256 | 16a3 | 2ec4 | bwtekus6.spt | bwtekusb2.inf | bwtekusb.dll |
| BTC262E-512 | 16a3 | 2ec4 | bwtekus6.spt | bwtekusb2.inf | bwtekusb.dll |
| BTC262P-512 | 16a3 | 2ecf | bwteku16.spt | bwtekusb2.inf | bwtekusb.dll |
| BTC263E-256 | 16a3 | 2ece | bwteku15.spt | bwtekusb2.inf | bwtekusb.dll |
| BTC264P-256 | 16a3 | 2ec3 | bwteku16.spt | bwtekusb2.inf | bwtekusb.dll |
| BTC264P-512 | 16a3 | 2ec3 | bwteku16.spt | bwtekusb2.inf | bwtekusb.dll |
| BTC264P-1024 | 16a3 | 2ec3 | bwteku16.spt | bwtekusb2.inf | bwtekusb.dll |
| BTC611E-512 | 16a3 | 2ec2 | bwtekus4.spt | bwtekusb2.inf | bwtekusb.dll |
| BTC611E-1024 | 16a3 | 2ec2 | bwtekus4.spt | bwtekusb2.inf | bwtekusb.dll |
| BTC613E-512 | 16a3 | 2ec2 | bwtekus4.spt | bwtekusb2.inf | bwtekusb.dll |
| BTC613E-1024 | 16a3 | 2ec2 | bwtekus4.spt | bwtekusb2.inf | bwtekusb.dll |
| BTC621E-512 | 16a3 | 2ec2 | bwtekus4.spt | bwtekusb2.inf | bwtekusb.dll |
| BTC621E-1024 | 16a3 | 2ec2 | bwtekus4.spt | bwtekusb2.inf | bwtekusb.dll |
| BTC651E-512 | 16a3 | 2ec2 | bwtekus4.spt | bwtekusb2.inf | bwtekusb.dll |
| BTC651E-1024 | 16a3 | 2ec2 | bwtekus4.spt | bwtekusb2.inf | bwtekusb.dll |
| BTC655E | 16a3 | 2ed3 | BTC655.spt | bwtekusb3.inf | bwtekusb.dll |
| BTC661E | 16a3 | 2ec2 | bwtekus4.spt | bwtekusb2.inf | bwtekusb.dll |
| BTC665E | 16a3 | 2ed1 | BTC665.spt | bwtekusb3.inf | bwtekusb.dll |

# Appendix D: Spectrometer’s USB Drivers File Destinations

| Items | Target Location |
| --- | --- |
| INF File bwtekusb2.inf bwtekusb3.inf | C:\windows\inf\ |
| CAT File fx2lp.cat cyusb3.cat | C:\windows\inf\ |
| Spt File bwtekusb.spt bwtekusb2.spt bwtekus4.spt bwtekus6.spt bwteku11.spt bwteku14.spt bwteku15.spt bwteku16.spt bwtekLC.spt BRC115.spt BTC655.spt BTC665.spt BTC675.spt | C:\windows\system32\drivers\bwtek\ |
| SYS File Cyusb.sys Cyusb3.sys | C:\windows\system32\ C:\windows\system32\drivers\ |
| DLL File CyUSB.dll WdfCoInstaller01009.dll | C:\windows\system32\ C:\windows\system32\drivers\ |
| DLL File BWTEKUSB.DLL | C:\windows\system32\ + The folder of your executable file. |

# Appendix E: SDK Function Groups

| USB Models | Supported Functions: |
| --- | --- |
| BRC111A | • USB Basic Functions • Supplementary Functions • AUX Port Functions 1 |
| BRC112P | • USB Basic Functions • Supplementary Functions • AUX Port Functions 1 • bwtekSetTTLOut (FROM AUX PORT FUNCTIONS 3) • bwtekReadTemperature |
| BRC112E | • USB Basic Functions • Supplementary Functions • AUX Port Functions 1 • bwtekSetTTLOut (FROM AUX PORT FUNCTIONS 3) |
| BRC115E | • USB Basic Functions • Supplementary Functions • AUX Port Functions 3 • bwtekReadTemperature |
| BRC115U | • USB Basic Functions • Supplementary Functions • AUX Port Functions 3 • bwtekReadTemperature |
| BRC115V | • USB Basic Functions • Supplementary Functions • AUX Port Functions 3 • bwtekReadTemperature |
| BRC642E-2048 | • USB Basic Functions • Supplementary Functions • AUX Port Functions 1 • bwtekSetTTLOut (FROM AUX PORT FUNCTIONS 3) |
| BRC711E-512 | • USB Basic Functions • Supplementary Functions • AUX Port Functions 1 |
| BRC711E-1024 | • USB Basic Functions • Supplementary Functions • AUX Port Functions 1 |
| BRC741E-512 | • USB Basic Functions • Supplementary Functions • AUX Port Functions 1 |

| USB Models | Supported Functions: |
| --- | --- |
| BRC741E-1024 | • USB Basic Functions • Supplementary Functions • AUX Port Functions 1 |
| BTC111E | • USB Basic Functions • Supplementary Functions • AUX Port Functions 1 |
| BTC112E | • USB Basic Functions • Supplementary Functions • AUX Port Functions 1 |
| BTC162E | • USB Basic Functions • Supplementary Functions • AUX Port Functions 1 |
| BTC261E-256 | • USB Basic Functions • Supplementary Functions • BTC261E & BTC262E Functions |
| BTC261E-512 | • USB Basic Functions • Supplementary Functions • BTC261E & BTC262E Functions |
| BTC261E-1024 | • USB Basic Functions • Supplementary Functions • BTC261E & BTC262E Functions |
| BTC261P-512 | • USB Basic Functions • Supplementary Functions • AUX Port Functions 1 • bwtekSetTTLOut (FROM AUX PORT FUNCTIONS 3) • BTC261P & BTC262P & BTC264P Functions • bwtekReadTemperature |
| BTC262A-256 | • USB Basic Functions • Supplementary Functions • AUX Port Functions 1 • bwtekSetTTLOut (FROM AUX PORT FUNCTIONS 3) • BTC263E & BTC262A Functions |
| BTC262E-256 | • USB Basic Functions • Supplementary Functions • BTC261E & BTC262E Functions • bwtekReadTemperature |
| BTC262E-512 | • USB Basic Functions • Supplementary Functions • BTC261E & BTC262E Functions |

| USB Models | Supported Functions: |
| --- | --- |
| BTC262P-512 | • USB Basic Functions • Supplementary Functions • AUX Port Functions 1 • bwtekSetTTLOut (FROM AUX PORT FUNCTIONS 3) • BTC261P & BTC262P & BTC264P Functions • bwtekReadTemperature |
| BTC263E-256 | • USB Basic Functions • Supplementary Functions • AUX Port Functions 1 • bwtekSetTTLOut (FROM AUX PORT FUNCTIONS 3) • BTC263E & BTC262A Functions |
| BTC264P-256 | • USB Basic Functions • Supplementary Functions • AUX Port Functions 1 • bwtekSetTTLOut (FROM AUX PORT FUNCTIONS 3) • BTC261P & BTC262P & BTC264P Functions • bwtekReadTemperature |
| BTC264P-512 | • USB Basic Functions • Supplementary Functions • AUX Port Functions 1 • bwtekSetTTLOut (FROM AUX PORT FUNCTIONS 3) • BTC261P & BTC262P & BTC264P Functions • bwtekReadTemperature |
| BTC264P-1024 | • USB Basic Functions • Supplementary Functions • AUX Port Functions 1 • bwtekSetTTLOut (FROM AUX PORT FUNCTIONS 3) • BTC261P & BTC262P & BTC264P Functions • bwtekReadTemperature |
| BTC611E-512 | • USB Basic Functions • Supplementary Functions |
| BTC611E-1024 | • USB Basic Functions • Supplementary Functions |
| BTC613E-512 | • USB Basic Functions • Supplementary Functions • AUX Port Functions 1 • AUX Port Functions 2 |
| BTC613E-1024 | • USB Basic Functions • Supplementary Functions • AUX Port Functions 1 • AUX Port Functions 2 |
| USB Models | Supported Functions: |
| BTC621E-512 | • USB Basic Functions • Supplementary Functions • AUX Port Functions 1 • AUX Port Functions 2 |
| BTC621E-1024 | • USB Basic Functions • Supplementary Functions • AUX Port Functions 1 • AUX Port Functions 2 |
| BTC651E-512 | • USB Basic Functions • Supplementary Functions • AUX Port Functions 1 • AUX Port Functions 2 |
| BTC651E-1024 | • USB Basic Functions • Supplementary Functions • AUX Port Functions 1 • AUX Port Functions 2 |
| BTC655E | • USB Basic Functions • Supplementary Functions • AUX Port Functions 3 • bwtekReadTemperature |
| BTC661E | • USB Basic Functions • Supplementary Functions • AUX Port Functions 1 • AUX Port Functions 2 |

# Appendix F: Default Parameters for bwtekTESTUSB

| USB Models | Timing Mode | Pixel # nPixelNo | Input Mode nInputMode | Integration Time Offset | Int.Time Range nTimingMode | Int. time unit settings | Array Location Dummy Pixels |
| --- | --- | --- | --- | --- | --- | --- | --- |
| BRC111A | 1 | 2048 | 1 | NA | 9 – 65535ms | milliseconds (ms) | 13 - 30 |
| BRC112P | 1 | 2048 | 1 | NA | 1 – 65535ms | milliseconds (ms) | 13 - 30 |
| BRC112E | 1 | 2048 | 1 | NA | 1 – 65535ms | milliseconds (ms) | 13 - 30 |
| BRC115E | 1 | 2048 | 1 | NA | 1,000us – 2,100,000,000us | microseconds (us) | 13 - 30 |
| BRC115U | 1 | 2048 | 1 | NA | 1,000us – 2,100,000,000us | microseconds (us) | 13 - 30 |
| BRC115V | 1 | 2048 | 1 | NA | 1,000us – 2,100,000,000us | microseconds (us) | 13 - 30 |
| BRC642E-2048 | 10 | 2048 | 2 | 6ms | 7 - 65535 ms | milliseconds (ms) | 4 - 6 |
| BRC711E-512 | 3 | 512 | 2 | NA | 2 – 65535ms | milliseconds (ms) | NA |
| BRC711E-1024 | 3 | 1024 | 2 | NA | 3 – 65535ms | milliseconds (ms) | NA |
| BRC741E-512 | 3 | 512 | 2 | NA | 2 – 65535ms | milliseconds (ms) | NA |
| BRC741E-1024 | 3 | 1024 | 2 | NA | 3 – 65535ms | milliseconds (ms) | NA |
| BTC111E | 1 | 2048 | 1 | NA | 9 – 65535ms | milliseconds (ms) | 13 - 30 |
| BTC112E | 1 | 2048 | 1 | NA | 5 – 65535ms | milliseconds (ms) | 13 - 30 |
| BTC162E | 1 | 2048 | 1 | NA | 5 – 65535ms | milliseconds (ms) | 13 - 30 |
| BTC261E-256 | 3 | 256 | 14 | NA | 0.010 – 65535ms | milliseconds (ms) | NA |
| BTC261E-512 | 3 | 512 | 13 | NA | 0.010 – 65535ms | milliseconds (ms) | NA |
| BTC261E-1024 | 3 | 1024 | 13 | NA | 0.010 – 65535ms | milliseconds (ms) | NA |
| BTC261P-512 | 1 | 512 | 2 | NA | 200us – 2,100,000,000us | microseconds (us) | NA |
| BTC262A-256 | 1 | 256 | 2 | 80us | 0.250 – 65535ms | microseconds (us) milliseconds (ms) | NA |
| BTC262E-256 | 3 | 256 | 14 | NA | 0.010 - 65535ms | milliseconds (ms) | NA |
| BTC262E-512 | 3 | 512 | 13 | NA | 0.010 – 65535ms | milliseconds (ms) | NA |
| BTC262P-512 | 1 | 512 | 2 | NA | 200us – 2,100,000,000us | microseconds (us) | NA |
| BTC263E-256 | 1 | 256 | 2 | 80us | 0.250 – 65535ms | microseconds (us) milliseconds (ms) | NA |
| BTC264P-256 | 1 | 256 | 2 | NA | 200us – 2,100,000,000us | microseconds (us) | NA |
| BTC264P-512 | 1 | 512 | 2 | NA | 200us – 2,100,000,000us | microseconds (us) | NA |
| BTC264P-1024 | 1 | 1024 | 2 | NA | 200us – 2,100,000,000us | microseconds (us) | NA |
| BTC611E-512 | 6 | 512 | 2 | 26ms | 27 – 65535ms | milliseconds (ms) | 4 - 9 |
| BTC611E-1024 | 6 | 1024 | 2 | 49ms | 50 – 65535ms | milliseconds (ms) | 4 - 9 |
| USB Models | Timing Mode | Pixel # nPixelNo | Input Mode nInputMode | Integration Time Offset | Int.Time Range nTimingMode | Int. time unit settings | Array Location Dummy Pixels |
| BTC613E-512 | 6 | 512 | 2 | 26ms | 27 – 65535ms | milliseconds (ms) | 4 - 9 |
| BTC613E-1024 | 6 | 1024 | 2 | 49ms | 50 – 65535ms | milliseconds (ms) | 4 - 9 |
| BTC621E-512 | 6 | 512 | 2 | 26ms | 27 – 65535ms | milliseconds (ms) | 4 - 9 |
| BTC621E-1024 | 6 | 1024 | 2 | 49ms | 50 – 65535ms | milliseconds (ms) | 4 - 9 |
| BTC651E-512 | 6 | 512 | 2 | 26ms | 27 – 65535ms | milliseconds (ms) | 4 - 9 |
| BTC651E-1024 | 6 | 1024 | 2 | 49ms | 50 – 65535ms | milliseconds (ms) | 4 - 9 |
| BTC655E | 1 | 2048 | 2 | 5601us | 6000us – 2,100,000,000us | microseconds (us) | 4 - 9 |
| BTC661E | 6 | 2048 | 2 | 20ms | 21 – 65535ms | milliseconds (ms) | 4 - 9 |
| BTC665E | 1 | 2048 | 2 | 5601us | 6000us – 2,100,000,000us | microseconds (us) | 4 - 9 |

# Appendix G: TTL Output

| USB Models | Multi-Porpose Pin | Shutter 1 & TTL Out 1 | Shutter 2 & TTL Out 2 | TTL Out 3 | TTL Out 4 |
| --- | --- | --- | --- | --- | --- |
| BRC111A | bwtekSetExtLaser bwtekSetExtShutter bwtekSetExtSync bwtekSetExtPulse | NA | NA | NA | NA |
| BRC112P | bwtekSetExtLaser bwtekSetExtShutter bwtekSetExtSync bwtekSetExtPulse | BwtekSetTTLOut (0,SetValue,Inverse,channel) | BwtekSetTTLOut (1,SetValue,Inverse,channel) | NA | NA |
| BRC112E | bwtekSetExtLaser bwtekSetExtShutter bwtekSetExtSync bwtekSetExtPulse |  |  |  |  |
| BRC115E |  | BwtekSetTTLOut (0,SetValue,Inverse,channel) | BwtekSetTTLOut (1,SetValue,Inverse,channel) | BwtekSetTTLOut (2,SetValue,Inverse,channel) | BwtekSetTTLOut (3,SetValue,Inverse,channel) |
| BRC115U |  | BwtekSetTTLOut (0,SetValue,Inverse,channel) | BwtekSetTTLOut (1,SetValue,Inverse,channel) | BwtekSetTTLOut (2,SetValue,Inverse,channel) | BwtekSetTTLOut (3,SetValue,Inverse,channel) |
| BRC115V |  | BwtekSetTTLOut (0,SetValue,Inverse,channel) | BwtekSetTTLOut (1,SetValue,Inverse,channel) | BwtekSetTTLOut (2,SetValue,Inverse,channel) | BwtekSetTTLOut (3,SetValue,Inverse,channel) |
| BRC642E-2048 | bwtekSetExtLaser bwtekSetExtShutter bwtekSetExtSync bwtekSetExtPulse |  |  |  |  |
| BRC711E-512 | bwtekSetExtLaser bwtekSetExtShutter bwtekSetExtSync bwtekSetExtPulse |  |  |  |  |
| BRC711E-1024 | bwtekSetExtLaser bwtekSetExtShutter bwtekSetExtSync bwtekSetExtPulse |  |  |  |  |
| BRC741E-512 | bwtekSetExtLaser bwtekSetExtShutter bwtekSetExtSync bwtekSetExtPulse |  |  |  |  |
| BRC741E-1024 | bwtekSetExtLaser bwtekSetExtShutter bwtekSetExtSync bwtekSetExtPulse |  |  |  |  |

| USB Models | Multi-Porpose Pin | Shutter 1 & TTL Out 1 | Shutter 2 & TTL Out 2 | TTL Out 3 | TTL Out 4 |
| --- | --- | --- | --- | --- | --- |
| BTC111E | bwtekSetExtLaser bwtekSetExtShutter bwtekSetExtSync bwtekSetExtPulse |  |  |  |  |
| BTC112E | bwtekSetExtLaser bwtekSetExtShutter bwtekSetExtSync bwtekSetExtPulse |  |  |  |  |
| BTC162E | bwtekSetExtLaser bwtekSetExtShutter bwtekSetExtSync bwtekSetExtPulse |  |  |  |  |
| BTC261E-256 | NA |  |  |  |  |
| BTC261E-512 | NA |  |  |  |  |
| BTC261E-1024 | NA |  |  |  |  |
| BTC261P-512 | bwtekSetExtLaser bwtekSetExtShutter bwtekSetExtSync bwtekSetExtPulse | BwtekSetTTLOut (0,SetValue,Inverse,channel) | BwtekSetTTLOut (1,SetValue,Inverse,channel) |  |  |
| BTC262A-256 | bwtekSetExtLaser bwtekSetExtShutter bwtekSetExtSync bwtekSetExtPulse | BwtekSetTTLOut (0,SetValue,Inverse,channel) | BwtekSetTTLOut (1,SetValue,Inverse,channel) |  |  |
| BTC262E-256 | bwtekSetExtLaser bwtekSetExtShutter bwtekSetExtSync bwtekSetExtPulse |  |  |  |  |
| BTC262E-512 | bwtekSetExtLaser bwtekSetExtShutter bwtekSetExtSync bwtekSetExtPulse |  |  |  |  |
| BTC262P-512 | bwtekSetExtLaser bwtekSetExtShutter bwtekSetExtSync bwtekSetExtPulse | BwtekSetTTLOut (0,SetValue,Inverse,channel) | BwtekSetTTLOut (1,SetValue,Inverse,channel) |  |  |
| BTC263E-256 | bwtekSetExtLaser bwtekSetExtShutter bwtekSetExtSync bwtekSetExtPulse | BwtekSetTTLOut (0,SetValue,Inverse,channel) | BwtekSetTTLOut (1,SetValue,Inverse,channel) |  |  |

| USB Models | Multi-Porpose Pin | Shutter 1 & TTL Out 1 | Shutter 2 & TTL Out 2 | TTL Out 3 | TTL Out 4 |
| --- | --- | --- | --- | --- | --- |
| BTC264P-256 | bwtekSetExtLaser bwtekSetExtShutter bwtekSetExtSync bwtekSetExtPulse | BwtekSetTTLOut (0,SetValue,Inverse,channel) | BwtekSetTTLOut (1,SetValue,Inverse,channel) |  |  |
| BTC264P-512 | bwtekSetExtLaser bwtekSetExtShutter bwtekSetExtSync bwtekSetExtPulse | BwtekSetTTLOut (0,SetValue,Inverse,channel) | BwtekSetTTLOut (1,SetValue,Inverse,channel) |  |  |
| BTC264P-1024 | bwtekSetExtLaser bwtekSetExtShutter bwtekSetExtSync bwtekSetExtPulse | BwtekSetTTLOut (0,SetValue,Inverse,channel) | BwtekSetTTLOut (1,SetValue,Inverse,channel) |  |  |
| BTC611E-512 | NA |  |  |  |  |
| BTC611E-1024 | NA |  |  |  |  |
| BTC613E-512 | bwtekSetExtLaser bwtekSetExtShutter bwtekSetExtSync bwtekSetExtPulse |  |  |  |  |
| BTC613E-1024 | bwtekSetExtLaser bwtekSetExtShutter bwtekSetExtSync bwtekSetExtPulse |  |  |  |  |
| BTC621E-512 | bwtekSetExtLaser bwtekSetExtShutter bwtekSetExtSync bwtekSetExtPulse |  |  |  |  |
| BTC621E-1024 | bwtekSetExtLaser bwtekSetExtShutter bwtekSetExtSync bwtekSetExtPulse |  |  |  |  |
| BTC651E-512 | bwtekSetExtLaser bwtekSetExtShutter bwtekSetExtSync bwtekSetExtPulse |  |  |  |  |
| BTC651E-1024 | bwtekSetExtLaser bwtekSetExtShutter bwtekSetExtSync bwtekSetExtPulse |  |  |  |  |
| BTC655E |  | BwtekSetTTLOut (0,SetValue,Inverse,channel) | BwtekSetTTLOut (1,SetValue,Inverse,channel) | BwtekSetTTLOut (2,SetValue,Inverse,channel) | BwtekSetTTLOut (3,SetValue,Inverse,channel) |

| USB Models | Multi-Porpose Pin | Shutter 1 & TTL Out 1 | Shutter 2 & TTL Out 2 | TTL Out 3 | TTL Out 4 |
| --- | --- | --- | --- | --- | --- |
| BTC661E | bwtekSetExtLaser bwtekSetExtShutter bwtekSetExtSync bwtekSetExtPulse |  |  |  |  |

# Appendix H: TTL Input

| USB Models | TTL In 1 | TTL In 2 |
| --- | --- | --- |
| BRC111A | bwtekGetExtStatus (channel) |  |
| BRC112P | BwtekGetTTLIn (0,*value,channel) | BwtekGetTTLIn (1,*value,channel) |
| BRC112E | bwtekGetExtStatus (channel) |  |
| BRC115E | BwtekGetTTLIn (0,*value,channel) | BwtekGetTTLIn (1,*value,channel) |
| BRC115U | BwtekGetTTLIn (0,*value,channel) | BwtekGetTTLIn (1,*value,channel) |
| BRC115V | BwtekGetTTLIn (0,*value,channel) | BwtekGetTTLIn (1,*value,channel) |
| BRC642E-2048 |  |  |
| BRC711E-512 | bwtekGetExtStatus (channel) |  |
| BRC711E-1024 | bwtekGetExtStatus (channel) |  |
| BRC741E-512 | bwtekGetExtStatus (channel) |  |
| BRC741E-1024 | bwtekGetExtStatus (channel) |  |
| BTC111E | bwtekGetExtStatus (channel) |  |
| BTC112E | bwtekGetExtStatus (channel) |  |
| BTC162E | bwtekGetExtStatus (channel) |  |
| BTC261E-256 |  |  |
| BTC261E-512 |  |  |
| BTC261E-1024 |  |  |
| BTC261P-512 | BwtekGetTTLIn (0,*value,channel) |  |
| BTC262A-256 | BwtekGetTTLIn (0,*value,channel) |  |
| BTC262E-256 |  |  |
| BTC262E-512 |  |  |
| BTC262P-512 | BwtekGetTTLIn (0,*value,channel) |  |
| USB Models | TTL In 1 | TTL In 2 |
| BTC263E-256 | BwtekGetTTLIn (0,*value,channel) |  |
| BTC264P-256 | BwtekGetTTLIn (0,*value,channel) |  |
| BTC264P-512 | BwtekGetTTLIn (0,*value,channel) |  |
| BTC264P-1024 | BwtekGetTTLIn (0,*value,channel) |  |
| BTC611E-512 |  |  |
| BTC611E-1024 |  |  |
| BTC613E-512 | bwtekGetExtStatus (channel) |  |
| BTC613E-1024 | bwtekGetExtStatus (channel) |  |
| BTC621E-512 | bwtekGetExtStatus (channel) |  |
| BTC621E-1024 | bwtekGetExtStatus (channel) |  |
| BTC651E-512 | bwtekGetExtStatus (channel) |  |
| BTC651E-1024 | bwtekGetExtStatus (channel) |  |
| BTC655E | BwtekGetTTLIn (0,*value,channel) | BwtekGetTTLIn (1,*value,channel) |
| BTC661E | bwtekGetExtStatus (channel) |  |

# Appendix I: Temperature Sensing

| USB Models | Temperature 1 | Temperature 2 |
| --- | --- | --- |
| BRC111A |  |  |
| BRC112P | bwtekReadTemperature (0x12,channel) |  |
| BRC112E |  |  |
| BRC115E | bwtekReadTemperature (0x0,channel) |  |
| BRC115U | bwtekReadTemperature (0x0,channel) |  |
| BRC115V | bwtekReadTemperature (0x0,channel) |  |
| BRC642E-2048 |  |  |
| BRC711E-512 |  |  |
| BRC711E-1024 |  |  |
| BRC741E-512 |  |  |
| BRC741E-1024 |  |  |
| BTC111E |  |  |
| BTC112E |  |  |
| BTC162E |  |  |
| BTC261E-256 | bwtekQueryTemperature (11,channel) |  |
| BTC261E-512 | bwtekQueryTemperature (11,channel) |  |
| BTC261E-1024 | bwtekQueryTemperature (11,channel) |  |
| BTC261P-512 | bwtekReadTemperature (0x0,channel) | bwtekReadTemperature (0x1,channel) |
| BTC262A-256 | bwtekReadTemperature (0x0,channel) | bwtekReadTemperature (0x1,channel) |
| BTC262E-256 | bwtekQueryTemperature (11,channel) |  |
| BTC262E-512 | bwtekQueryTemperature (11,channel) |  |
| BTC262P-512 | bwtekReadTemperature (0x0,channel) | bwtekReadTemperature (0x1,channel) |
| BTC263E-256 | bwtekReadTemperature (0x0,channel) | bwtekReadTemperature (0x1,channel) |
| BTC264P-256 | bwtekReadTemperature (0x0,channel) | bwtekReadTemperature (0x1,channel) |

| USB Models | Temperature 1 | Temperature 2 |
| --- | --- | --- |
| BTC264P-512 | bwtekReadTemperature (0x0,channel) | bwtekReadTemperature (0x1,channel) |
| BTC264P-1024 | bwtekReadTemperature (0x0,channel) | bwtekReadTemperature (0x1,channel) |
| BTC611E-512 |  |  |
| BTC611E-1024 |  |  |
| BTC613E-512 |  |  |
| BTC613E-1024 |  |  |
| BTC621E-512 |  |  |
| BTC621E-1024 |  |  |
| BTC651E-512 |  |  |
| BTC651E-1024 |  |  |
| BTC655E | bwtekReadTemperature (0x10,channel) | bwtekReadTemperature (0x11,channel) |
| BTC661E |  |  |

# Appendix J: ADC & DAC

| USB Models | ADC (Input) | DAC (Output) |
| --- | --- | --- |
| BRC111A |  |  |
| BRC112P |  |  |
| BRC112E |  |  |
| BRC115E | bwtekGetAnalogIn (0,channel) | bwtekSetAnalogOut (0,setvalue,channel) |
| BRC115U | bwtekGetAnalogIn (0,channel) | bwtekSetAnalogOut (0,setvalue,channel) |
| BRC115V | bwtekGetAnalogIn (0,channel) | bwtekSetAnalogOut (0,setvalue,channel) |
| BRC642E-2048 |  |  |
| BRC711E-512 |  |  |
| BRC711E-1024 |  |  |
| BRC741E-512 |  |  |
| BRC741E-1024 |  |  |
| BTC111E |  |  |
| BTC112E |  |  |
| BTC162E |  |  |
| BTC261E-256 |  |  |
| BTC261E-512 |  |  |
| BTC261E-1024 |  |  |
| BTC261P-512 | bwtekGetAnalogIn (0,channel) |  |
| BTC262A-256 | bwtekGetAnalogIn (0,channel) |  |
| BTC262E-256 |  |  |
| BTC262E-512 |  |  |
| BTC262P-512 | bwtekGetAnalogIn (0,channel) |  |
| BTC263E-256 | bwtekGetAnalogIn (0,channel) |  |
| BTC264P-256 | bwtekGetAnalogIn (0,channel) |  |
| BTC264P-512 | bwtekGetAnalogIn (0,channel) |  |
| BTC264P-1024 | bwtekGetAnalogIn (0,channel) |  |
| BTC611E-512 |  |  |
| BTC611E-1024 |  |  |
| BTC613E-512 |  |  |
| USB Models | ADC (Input) | DAC (Output) |
| BTC613E-1024 |  |  |
| BTC621E-512 |  |  |
| BTC621E-1024 |  |  |
| BTC651E-512 |  |  |
| BTC651E-1024 |  |  |
| BTC655E | bwtekGetAnalogIn (0,channel) | bwtekSetAnalogOut (0,setvalue,channel) |
| BTC661E |  |  |

# Appendix K: X-axis Reverse Settings

Some spectrometer may have their x-axis reversed for production purposes.

These models will first have their pixel values reversed.

For example if the physical location on the detector’s pixel array is Pixel 2047, the software will flip the detector array so pixel 2047 will be Pixel 0.

Our End User, BWSpec software, automatically accounts for this reversal.

In the para.ini (EEPROM) there is a line called “xaxis_data_reverse.”  When the value of xaxis_data_reverse=1” the unit will need its x-axis to be reversed.  When xaxis_data_reverse=0, no additional steps need to be performed.

| USB Models | X-axis Reversed | Comment |
| --- | --- | --- |
| BRC111A | No |  |
| BRC112P | Yes | xaxis_data_reverse=1 |
| BRC112E | BRC112E: No, BRC112U: No, BRC112V: Yes | xaxis_data_reverse=1 |
| BRC115E | No |  |
| BRC115U | No |  |
| BRC115V | No |  |
| BRC642E-2048 | Yes | xaxis_data_reverse=1 |
| BRC711E-512 | No |  |
| BRC711E-1024 | No |  |
| BRC741E-512 | Yes | xaxis_data_reverse=1 |
| BRC741E-1024 | Yes | xaxis_data_reverse=1 |
| BTC111E | No |  |
| BTC112E | No |  |
| BTC162E | Yes | xaxis_data_reverse=1 |
| BTC261E-256 | Yes | xaxis_data_reverse=1 |
| BTC261E-512 | Yes | xaxis_data_reverse=1 |
| BTC261E-1024 | Yes | xaxis_data_reverse=1 |
| BTC261P-512 | Yes | xaxis_data_reverse=1 |
| BTC262A-256 | Yes : For '+1 Order' configured spectrometers | xaxis_data_reverse=1 |
| BTC262E-256 | Yes | xaxis_data_reverse=1 |
| BTC262E-512 | Yes | xaxis_data_reverse=1 |
| BTC262P-512 | Yes | xaxis_data_reverse=1 |
| BTC263E-256 | Yes : For '+1 Order' configured spectrometers | xaxis_data_reverse=1 |
| BTC264P-256 | Yes | xaxis_data_reverse=1 |
| BTC264P-512 | Yes | xaxis_data_reverse=1 |
| BTC264P-1024 | Yes | xaxis_data_reverse=1 |
| BTC611E-512 | No |  |
| BTC611E-1024 | No |  |
| USB Models | X-axis Reversed | Comment |
| BTC613E-512 | No |  |
| BTC613E-1024 | No |  |
| BTC621E-512 | Yes | xaxis_data_reverse=1 |
| BTC621E-1024 | Yes | xaxis_data_reverse=1 |
| BTC651E-512 | No |  |
| BTC651E-1024 | No |  |
| BTC655E | No |  |
| BTC661E | No |  |
| BTC665E | Yes | xaxis_data_reverse=1 |

# Appendix L: Dynamically Assigned Channel Numbers

## Procedure:

- Call function InitDevices() ( Initialize the USB devices

- Call function bwtekSetupChannel() ( get channel array  (channel[]

- Call function GetDeviceCount() ( get connected USB spectrometers count

- Loop read all spectrometer’s channel, c.code and EEPROM.

## Sample Code:

The below sample code is an excerpt from the C#_2.1 sample code from Demo 2.1 located in the SDK-S v1.0.0.1 sample code folder.
Function calls from the SDK User Manual used below are in Red Font below.

public struct Spec_Para_Struct

{

public int usbtype;                //2=USB2.0, 3=USB3.0

public int channel;

public string cCode;

public string model;

public string spectrometer_name;

public int spectrometer_type;

public int pixel_number;

public int timing_mode;

public int input_mode;

public int xaxis_data_reverse;

public double inttime;

public int inttime_int;

public double inttime_min;

public int inttime_base;

public int inttime_unit;

public double coefficient_a0;

public double coefficient_a1;

public double coefficient_a2;

public double coefficient_a3;

public double coefficient_b0;

public double coefficient_b1;

public double coefficient_b2;

public double coefficient_b3;

}

public static Spec_Para_Struct[] spec_para = new Spec_Para_Struct[32];

bool retcode = InitDevices();	//Initialize USB Device

if (retcode)

{

byte[] channel = new byte[32];

int retcode1 = bwtekSetupChannel(-1, channel);	//Get all available channels to array

if (retcode1 > 0)

{

device_count = GetDeviceCount();

for (int i = 0; i < device_count; i++)	//Loop get channel, ccode and read eeprom

{

if (channel[i] < 32)

{

spec_para[i].channel = channel[i];	//Get channel

byte[] tmp_pccode = new byte[8];

int ret = GetCCode(tmp_pccode, i);

spec_para[i].cCode = ASCIIEncoding.ASCII.GetString(tmp_pccode); //Get CCode

int tmp_usbtype = 0;

ret = GetUSBType(ref tmp_usbtype, i);

spec_para[i].usbtype = tmp_usbtype; 	//Get USB type

string filename = System.IO.Path.GetDirectoryName(System.Windows.Forms.Application.ExecutablePath) + "\\para.ini";

ret = bwtekReadEEPROMUSB(filename, spec_para[i].channel);  //Read EEPROM to file para.ini

Load_Para(filename, i);		//Read spectrometer’s parameter from file.

}

}

}

}

# Text boxes

Parameter

Export