From: Rob Roberts <rob.roberts@bio-logic.us>

Subject: RE: SP-300 used as Bipotentiostat

Date: January 31, 2017 at 4:49:47 AM PST

To: "Brunschwig, Bruce S." <bsb@caltech.edu>, Bill Eggers <bill.eggers@bio-logic.us>, Arnold Forman <arnold.forman@bio-logic.us>

Cc: "Nunez, Paul D." <pnunez@caltech.edu>, "Fermin, David J." <davidcal@caltech.edu>

Bruce,

To run the SP-300 as a bi-stat, it needs to run in CE-to-ground mode which is a selection in the Advanced Settings, and special lead connections.  See attached photos of SP-300 connected to RRDE system (two WE, with a common RE and CE).  In this mode, the ground leads become the P2 (counter) leads, and the regular P2 (dark blue) leads are not used.

Best regards,

Rob Roberts

Strategic Marketing and Sales Manager

Bio-Logic USA

9050 Executive Park Drive

Suite 110C

Knoxville,TN 37923

Tel: 865-769-3800

Email: rob.roberts@bio-logic.us

[www.ec-lab.com](http://www.ec-lab.com/)

