# From: Rob Roberts <rob.roberts@bio-logic.us>

Subject: RE: SP-300 used as Bipotentiostat

Date: January 31, 2017 at 4:49:47 AM PST

To: "Brunschwig, Bruce S." <bsb@caltech.edu>, Bill Eggers <bill.eggers@bio-logic.us>, Arnold Forman <arnold.forman@bio-logic.us>

Cc: "Nunez, Paul D." <pnunez@caltech.edu>, "Fermin, David J." <davidcal@caltech.edu>

Bruce,

To run the SP-300 as a bi-stat, it needs to run in CE-to-ground mode which is a selection in the Advanced Settings, and special lead connections.  See attached photos of SP-300 connected to RRDE system (two WE, with a common RE and CE).  In this mode, the ground leads become the P2 (counter) leads, and the regular P2 (dark blue) leads are not used.

# Best regards,

Rob Roberts

Strategic Marketing and Sales Manager

Bio-Logic USA

9050 Executive Park Drive

Suite 110C

Knoxville,TN 37923

# Tel: 865-769-3800

Email: rob.roberts@bio-logic.us

www.ec-lab.com