Kratos UHV System Sample Loading

Preliminary

- Rinse and dry samples and stubs with low residue solvent (methanol, isopropanol (IPA), trichloroethylene, etc.) before loading into chamber. If you use any solvents other than IPA rinsh with IPA after the other solvents. The use of the ultrasonic cleaner is recommended. Also you can dry samples in the oven.

- Be sure samples are clean and dry – do not use the XPS chambers to clean your samples.

- If you are unsure about  your samples, discuss with a GLA.

- The Edwards turbo controller for the load lock is located on the floor below the load lock.

Sample Holders

- Pucks: a few small samples may be affixed to the pucks via the short screws.! Screws must not show below stub lip!!!

- Bars with screws: This bar allows for the loading of several samples affixed very well by screws.

- Spring loaded bar: The spring loaded bar allows for the loading of up to 4 samples with constant height. This bar is particularly good for angle resolved measurements.

- Pots: these are good for powder or nanoparticle samples. There is one special pot with a variable depth.  The depth is controlled by a screw in rod. This holder is good for bringing different volumes of powder or nanoparticle samples to surface level.

- TPD stages: These are assembled by the user prior to use and are specific for using the resistive heating stage in the SAC for doing TPD measurements.

Loading Samples – Load Lock

- Affixed your clean and dry sample to a sample holder.

- Verify the load-lock-to-STC valve is closed and system is normal, the pressure in the STC <8 × 10-9 Torr and SAC < 1 × 10-9 Torr. If this is not the case, discuss with the previous user or a GLA. Ensure there is enough space on the STC magazine for your sample holder. If necessary, remove a sample from the chamber before proceeding.

- Vent the load lock:

  - The  Edwards turbo controller (cart below Load Lock (LL) ) should displays the LL pressure. If it is not, press the menu button ( ) until the pressure is displayed.

  - Highlight ‘Turbo ON Back ON’ and push enter ()This brings you to the ‘System Control’ menu.

  - Highlight ‘System ON’ and press enter again to turn off the pump. Unlock the clamps holding the load lock closed.

- Wait ~90 s for the chamber to vent then open by pulling back gently on the load lock cylinder.

- Load Sample: The transfer claws have tiny pins that hold samples in place when moving the arm. Take care not to damage the pins. Load pucks and bars on the claw using the TOP groove.

- Check that the o-ring is fitted nicely in the groove on the load lock cylinder (left, load lock arm side), and close the load lock by sliding the cylinder shut.  Reclamp the cylinder using the right camp first.

- Be sure the turbo has spun down. On the Edwards turbo controll with ‘System OFF’ highlighted, push enter to start the turbo. Return to the pressure screen by pressing the menu button. Wait until pressure on turbo status window reads in the 10-6 Torr pressure range.

Loading Samples – STC

- Verify the load lock arm is fully retracted and the green light is lit in the sensor at the end of the arm.

- Verify the load-lock-to-STC and STC-to-SAC Flap valves are both closed.

- In the “Vacuum” panel of the “Instrument Manual Contol” window, verify the “Safe Manual” radio button is selected.

- Click on the small red rectangle, which represents the load-lock-to-STC valve, to bring up a green light bulb button. Click on the button and verify that the load-lock-to-STC valve opens.

- Carefully move the load lock transfer arm into the STC by turning the rotary drive on top of the load lock. Do not turn the rotary drive too quickly and do not apply force to the drive. Align the magazine by turning the rotary drive under the STC (near lamp controller). Place sample on the magazine, then move the magazine north or toward the back wall to move the sample holder off the load lock transfer arm hook.

- Fully retract the load lock transfer arm and close the load-lock-to-STC valve by clicking on the valve in the software and selecting the green light bulb that appears.

Loading Samples – Portable Sample Transfer Suitcase

The portable sample transfer suitcase is designed to allow transfer of samples from inert atmosphere directly into UHV without exposure to air. It is only compatible with sample pucks or pots. The suitcase arm loads onto the quickflange port on the south side of the STC facing the door. It is mounted with the gate valve oriented vertical, so the valve actuator mechanism is at 12 o’clock.

- Mount your sample on a puck under an inert or UHV atmosphere. Ensure the sample is secure, since the puck will be vertical during sample transport.

- With the arm’s gate valve open and the valve actuator oriented at 12 o’clock, move the claw out of the arm so it can be accessed.

Figure 1 Puck mounted on portable transport arm claw.

Load the puck using the bottom groove on the claw such that the sample faces up with the gate valve actuator mechanism at 12 o’clock. Be careful not to damage the pin at the end of the claw that holds the puck in place.

- Fully retract the claw back into the suitcase, Lock the magnet, and close the gate valve.

Figure 2 Valve to STC and Transport KF Flange.

Hold the suitcase arm such that the VAT label on the valve is oriented up while transferring the arm to the Kratos. In this orientation, the sample is vertical, but the claw holds the puck from beneath. Holding it the wrong way may cause the sample to fall out during transport. Use the screw on the end of the arm to lock the magnet in place during the transfer.

- At the Kratos, check that the pressure in the STC is < 10-7 Torr and that the Load Lock Turbo is running. If not start it up and wait till the load lock is at < 10-7-Torr.

- Check that the valve to the STC is closed by turning the valve clockwise (the one attached to the KF Flange that faces down), Figure 2. TAKE CARE: there are no interlocks on the valves for the transport arm. You can vent the STC by mistake!!! If the valve to the STC is open, close it, and call a GLA.

- Vent the load lock as described above.

- If the valve to the Transport is closed, open it, Figure 3. This valve should be left open all the times to prevent forming pressure differentials that can damage the pump or compromise the atmosphere of the STC if valves are inappropriately handled.

- Attaching the Transport arm to the Transport KF Flange, Figure 2:

  - Remove the KF cap from STC/transport arm port and the cap on transport arm.

  - Attach the transport arm (transport arm valve facing up) to STC, such that that the valve actuator is at 12 o’clock.

Figure 3 Valve to Transport and Valve to Load Lock from Turbo pump

Fasten the arm with the clamp for the KF with the bolt/screw facing down, Figure 2.

- Restart the load lock turbo pump and check that the “V to transport” valve is open, Figure 3, so the space between the transport arm and the STC gets pumped down.

- Put sign on the Load Lock turbo that the transport are is attached and the operation of the load lock turbo should  not be changed.

- Keep the valve on PTA closed and wait until the turbo pump reaches full speed and the pressure is <10-4 Torr

- Vent load lock again (venting load lock is only possible if turbo pump is up to speed).

- As soon as pressure gauge on load lock reads 7.5x102 Torr (90 seconds), i.e., the load lock is still under a slight vacuum, start the load lock turbo pump again and open the hand valve on PTÅ, Figure 4, the handle facing up, NOT THE VALVE to STC FACING DOWN!!!!)

- Wait until load lock pressure is <10-6 Torr, takes about 1h)

- Open the regular load lock/STC gate valve using the computer to check if it can actually be opened and no interlock prevents it (if you cannot open this valve, check if load lock is at low enough pressure and the normal load lock transport arm is fully retracted; if these were not the problem, stop here and call a GLA)

- CRITICAL STEP: Check that

  - The regular load lock/STC gate valve is open

  - “V to transport valve” is open

Figure 4 Transport arm attached to the KF Flange with transport arm valve handle facing up.

Hand valve on transport arm (the one facing up) is open

  - ONLY THEN it is safe to open the valve to STC (the one facing down), Figure 2. If a, b or c is not open STOP HERE and restart from step (6).

- Unlock the magnet and extend the transport arm into the STC chamber.

- Transport your sample from the transport arm claw onto the claw for the STC/SAC arm. This may require gently manipulating the suitcase arm to adjust its height inside the STC.

- Fully retract Transport arm back into its chamber, lock the magnet, and close the PTA valve.

- Close valve to STC (the one facing down, Figure 2), make sure that it is fully closed.

- Close the regular load lock STC gate via the computer.

- Check that the pressure in STC goes back to normal (if not the manual valve might still be open, close it.)

- Vent the load lock as described above and remove the suitcase arm. This is important to prevent others from accidentally damaging the arm assembly. Place the arm on the sample prep counter with the valve closed and the plastic cover placed over the end.

- Wipe the quickflange blank with a low residue solvent, place back on the suitcase arm port, and secure.

Loading Samples – SAC

- Verify that the pressure in the STC <1× 10-8 Torr and in the SAC <3× 10-9 Torr.

- Verify the STC-to-SAC arm is fully retracted and the green light is lit in the sensor at the end of the arm.

- In the “Vacuum” panel of the “Instrument Manual Contol” window, verify the “Safe Manual” radio button is selected.

- Click on the large red rectangle, which represents the STC-to-SAC valve, to bring up a green light bulb button. Click on the button and verify that the STC-to-SAC valve opens. This will take 10 – 15 seconds.

- The stage can be adjusted using the buttons in the “Stage” panel of the “Instrument Manual Control” window or using the “Autostage Manual Control” device. The top sample rotation section is not used with the sample holders we currently have. Pressing the button in the desired direction and then pressing the opposite direction allows for faster motion of the stage.

- Set the stage in the sample load position. To do this, load the position table /C=/stage positions/Stub_bar_load_unload.dset in the “Stage” section of the “Instrument Manual Control” window. Then, select the “Stub Load” position and click “Go to.” Wait until the stage stops moving.

- Load the desired sample from the magazine onto the STC-to-SAC arm. Move the magazine north toward the back wall to hook the sample holder on the arm. The arm can then be directed past the magazine by aligning the magazine such that the arm can move through the center or past the right end (when looking towards the SAC) of the magazine.

- Bring the STC-to-SAC arm into the SAC and load the sample holder onto the stage. The stage can hold up to two pucks or one bar.

- For bars, it will be necessary to drive the stage to the right in order to view the location where the bar is held onto the STC-to-SAC arm.

- Manually move the stage south towards the door to unhook the claw from the sample holder. The arm can be gently wiggled to prevent it from sticking.

- Withdraw the arm fully into the STC and close the STC-to-SAC valve by clicking on the valve in the software and selecting the green light bulb that appears.

Unloading Samples

- Verify the system is normal and all data collection or processes have finished. The “Safe Manual” radio button should be selected in the “Vacuum” panel in the “Instrument Manual Control” window.

- Open the STC-to-SAC valve by clicking on the large red rectange and selecting the green light bulb that appears. Wait 10-15 seconds for the valve to open.

- Set the stage in the sample unload position. To do this, load the position table /C=/stage positions/Stub_bar_load_unload.dset in the “Stage” panel of the “Instrument Manual Control” window. Then, select the “Stub Unoad” position and click “Go to.” Wait until the stage stops moving. If unloading a bar, drive the stage to the right until the end of the bar is visible in the window.

- Move the STC-to-SAC arm into the SAC, taking care to avoid the magazine. Slowly move the arm past the sample holder until the claw is lined up with the sample holder.

- Manually move the stage toward the back of the room using the “Autostage Manual Control” device. The arm can be gently wiggled to allow for easier loading.

- Move the sample into the STC and place the sample on the magazine and move the magazine south toward the door to unhook the arm.

- Fully withdraw the STC-to-SAC arm and close the STC-to-SAC valve by clicking on the valve in the software and selecting the green light bulb that appears.

- Verify that the load lock pressure is in the 10-6 Torr range.

- Open the load-lock-to-STC gate valve by clicking on the samall red rectangle in the “Vacuum” panel of the “Instrument Manual Control” window and selecting the green light bulb that appears.

- Move the load lock arm into the STC by rotating the rotary drive above the load lock clockwise slowly and carefully.

- Hook the sample holder onto the load-lock-to-STC arm by moving the magazine south toward the door.

- Withdraw the load-lock-to-STC arm by turning the rotary drive counterclockwise carefully and slowly. Watch the arm move through the window on the load lock. Once the claw portion of the arm disappears in the window, it is less than a full turn until the arm is fully retracted. Watch for the sensor on the end of the arm to turn green. Be careful not to jam the arm.

- Close the load-lock-to-STC valve by clicking on the valve in the “Vacuum” panel of the “Instrument Manual Control” window and selecting the green light bulb that appears.

- Vent the load lock as described above. Once vented, remove the sample and pump the load lock back down as described above.