Shaun Hodgetts shawn@qdusa.co

fix of transport:

- Turn off power on qd and computer

- wait 30sec and turn on

- Write all values of the calibration files down (see below) f7:sample transport

- revame svomacro.cfg to svomacro.alg

- delete all cfg files

- rename svomacro.alg svomacro.cfg

- restart mpmsr2 and hit space bar to get out of error dialoges about missing files etc

- go to sample transport and turn power off

- move transport by hand to trip lower limit switch and position transport in middle

- turn power on

- go to f2 and initialize transport

Parameters from new squid computer

Calibration Parameres

| Calibration parameters |  |
| --- | --- |
| Magu | 1532.310 |
| Long | -8151.969 |
| Trans | 7500.000 |
| Field Den | 5.000 |
| Long | 307.300 |
| Trans | 300.000 |
| long Reg | 1.825 |
| Trans Reg | 31.900 |
| Min | 8 |
| T Rat | 96.000 |

| Squid Settle |  |
| --- | --- |
| None | 0.050 |
| 100 Hz | 0.050 |
| 10 Hz | 0.080 |
| 1Hz | 0.800 |
| Ellip | 0.800 |

| Temp Calib |  |
| --- | --- |
| Temp Tol | 0.005 |
| Low Temp | 4.430 |
| fill temp | 4.150 |
| fill time | 650.000 |
| fill power | 3.500 |
| oven factor | 15.000 |
| Max T | 800.000 |
| min warming | 5.000 |
| Therm Consis | 3.000 |
| Low T | 2.000 |

| Low Field Cal |  |
| --- | --- |
| Coil Cal1 | 82.200 |
| coil cal 2 | 45.000 |
| fluxgate | 10.000 |
| DACrange | 0.1000 |
| Remanent Tol | 0.0050 |
| Mag Warm | 5.000 |
| Mag Cool | 30.000 |
| Quench warm | 30.00 |
| quench cool | 180.000 |
| coil field wait | 1.000 |
| field mult | 0.000 |