## Sheet1
Lamp Serial # | 7~1803 | Newport Corp Stratford, CT
2008-06-06
Lamp Data File Name: | Q4300.DAT
Wavelength File: | Wavelengths.dat
Standard 1 Data File: | clonedstd.mat | REPORT OF CALIBRATION
7~1535 | 7~1535 | of One Standard of Spectral Irradiance (250 - 2400 nm)
Standard 2 Data File: | clonedstd.mat
7~1536 | 7~1536 | Oriel Part# 63358 Lamp Serial Number: 7~1803
Calibration Date: | 2008-06-06
1. Material
One 45 watt, quartz halogen, tungsten filament lamp with designation 7~1803
Fitting Coefficients:
A | 41.7915 | 2. Calibration Procedures
B | -4838.9814
C | 0.77864 | The lamp was pre-treated and calibrated using custom designed equipment in Oriel's calibration facility and modeled on the equipment and procedures described in the NBS Special Publication 250-20.
D | 543.4609
E | -426311.7821 | 3. Calibration Reference
F | 127392931.1747
G | -13902811647.7036 | The working standards used as reference were 7~1535 and 7~1536 which, in turn, were calibrated using F766 standard traceble to NIST reference standard: F-525.
H | 0
4. Results
Lamp Voltage Start | 109.7
Lamp Voltage: | 7 | The spectral irradiance produced by Standard 7~1803 at a distance of 0.5 m when operated at the calibration current is provided in a number of formats to facilitate the use of this tool. Irradiance levels at wavelengths for which NIST provides data,
are provided for best traceability to NIST. A fitting formula is provided, found using NIST suggested method, slightly modified for faster convergence when using commercial algorithms, is listed with the data. Irradiance values, calculated using this formula, are tabulated on the attached sheets in 10 nm increments and are included in an ASCII format in 1 nm increments on the PC format diskette included with this product.
Calibrated by:Bob Fusaris
Date:
2008-06-06
OPERATION OF THE IRRADIANCE STANDARD
Safety Considerations
There are three hazards associated with operation of this lamp.
1. Ultraviolet Radiation
2. Heat
3. Electrical
1. Ultraviolet radiation
These lamps emit low levels of ultraviolet radiation, UVA, UVB and UVC. Since these are potentially hazardous to the eyes and the skin, we recommend that you wear UV goggles and limit exposure of the skin. You will find a discussion on ultraviolet safety on page 1-191 of Oriel Volume II catalog. This page also shows the current recommended Threshold Limit Values for occupational exposure to ultraviolet. You should use the irradiance data and the radiation square law to estimate whether UV will be a problem. Obviously it depends how close you work to the lamp, how long you operate the lamp etc. UV goggles and glasses are available from Oriel. To limit skin exposure in an open laboratory environment we recommend you fabricate shields to enclose the lamp. These should be large enough (at least 0.5 m from the lamp) to ensure the lamp is effectively in the open air. Small enclosures will change the thermal and optical operating conditions of the lamp and hence the calibration.
2. Heat
The surface of the lamp is hot enough to burn skin or ignite combustible materials. Additionally, dark combustible materials close to the lamp may ignite.
3. Electrical hazard
The lamp operates at 6.5 A and approximately 7 V. Under certain conditions these are considered hazardous current and voltage levels and we recommend that you post appropriate warning signs since "live terminals" can be exposed.
Mounting the lamp
The irradiance standard is intended for vertical operation in the open air. The lamp should be held in an Oriel Model 63365 Lamp Mount or similar. All calibration and life data available refers to this form of operation. Different types of operation, for example horizontal operation or operation in an enclose will result in different irradiation and may shorten the useful life.
Power source
The lamps require a source of well regulated constant current such as the Oriel 68831 Radiometric Power Supply. The current source needs to be accurately calibrated to reproduce the calibration current. Deviations from the calibration value result in changes in the irradiance produced with the highest discrepancies in the ultraviolet region of the spectrum.
Lamp life
Studies at NIST of the 1kW type of tungsten halogen show different drift rates with time for lamps tested. While the drift (in % per 100 operating hours) was worst at 250 nm, the values (at 250 nm) for the lamps tested varied from 0.5 % to 4.8 %. We reproduce the NIST data for your information.
Lamp # 7~1803
Wavelength Lamp#
(nm) 1 2 3 4
250 0.50 4.3 4.0 4.8
350 0.45 3.4 2.9 3.35
450 0.40 2.90 2.30 2.55
555 0.40 2.55 1.90 2.00
654.6 0.40 2.35 1.60 1.65
800 0.40 2.1 1.35 1.3
1300 0.35 1.7 0.8 0.65
1600 0.35 1.6 0.75 0.5
Please note that these drift rates are for constant operation of the lamp. We expect worse data when the lamp is run for short periods as the filament is subjected to thermal shock each time the lamp is started. Oriel DC power supplies for these lamps have a "soft start" and “soft shutdown” features to mitigate this problem. Gradual increase to the operating current over a period of about a minute should allow you to approach the drift rates above when you operate the standard lamp for shortened periods.
It is recommended that the lamp be recalibrated after 50 hours of use.
NIST Results
The principal set of results is provided at NIST specified wavelength data points. Estimated Transfer Uncertainty (2δ) is as follows. The NIST Uncertainty of Reported Values for the F-525 standard lamp with respect to SI units are listed for information.
Wavelength (nm): 250 350 654.6 900 1300 1600 2000 2400
NIST Uncertainty 2.23 1.35 1.01 1.34 1.42 1.89 3.29 6.50
F766 Uncertainty 2.00 2.00 2.00 2.00 2.00 2.00 2.00 2.00
Oriel Transfer
Uncertainty 2.0 1.5 1.5 1.5 1.5 2.0 2.0 2.0
Total (SQRT of Quadrature sum) 3.6 2.8 2.7 2.8 2.9 3.4 4.3 7.1
Fig. 1 Reference points for the 50 cm (500 mm) calibration distance
CALIBRATION RESULTS
Model No. 63358 Quartz Tungsten Halogen lamp
Lamp Serial No.: 7~1803
Spectral Range: 250-2400nm
Lamp Current: 6.50 Amps
Voltage (ref. Only) = 7 V
Room Temperature: 25 C
Fitting equation for non-NIST wavelengths (divide values by 10 to obtain results in μW/cm2 nm):
Irradiance (mW/m2 nm) = λ -5 * exp(A + B/λ) * (C + D/λ + E/λ2 + F/λ3 + G/λ4 + H/λ5)
Where:
A= 41.7915090488591
B=-4838.98137717923
C=0.778641040863826
D=543.460937982459
E=-426311.782126541
F=127392931.174736
G=-13902811647.7036
H=0

## Sheet2
Irradiance data for lamp# 7~1803
nm | (mW/m2nm) | nm | (mW/m2nm) | nm | (mW/m2nm) | nm | (mW/m2nm) | nm | (mW/m2nm)
280 | 0.02109 | 291.1 | 0.034364 | 302.2 | 0.053413 | 313.3 | 0.079725 | 324.4 | 0.11488
335.5 | 0.16049 | 346.6 | 0.21815 | 357.7 | 0.28938 | 368.8 | 0.37556 | 379.9 | 0.47788
391 | 0.59733 | 402.1 | 0.73464 | 413.2 | 0.89026 | 424.3 | 1.0644 | 435.4 | 1.2569
446.5 | 1.4675 | 457.6 | 1.6954 | 468.7 | 1.94 | 479.8 | 2.2 | 490.9 | 2.4744
502 | 2.7616 | 513.1 | 3.0603 | 524.2 | 3.3688 | 535.3 | 3.6855 | 546.4 | 4.0087
557.5 | 4.3369 | 568.6 | 4.6683 | 579.7 | 5.0013 | 590.8 | 5.3344 | 601.9 | 5.6661
613 | 5.995 | 624.1 | 6.3198 | 635.2 | 6.6393 | 646.3 | 6.9523 | 657.4 | 7.2578
668.5 | 7.555 | 679.6 | 7.8429 | 690.7 | 8.121 | 701.8 | 8.3885 | 712.9 | 8.645
724 | 8.8901 | 735.1 | 9.1233 | 746.2 | 9.3445 | 757.3 | 9.5534 | 768.4 | 9.7499
779.5 | 9.934 | 790.6 | 10.1056 | 801.7 | 10.2649 | 812.8 | 10.4119 | 823.9 | 10.5467
835 | 10.6696 | 846.1 | 10.7808 | 857.2 | 10.8805 | 868.3 | 10.969 | 879.4 | 11.0466
890.5 | 11.1135 | 901.6 | 11.1703 | 912.7 | 11.217 | 923.8 | 11.2542 | 934.9 | 11.2822
946 | 11.3013 | 957.1 | 11.3119 | 968.2 | 11.3143 | 979.3 | 11.309 | 990.4 | 11.2963
1001.5 | 11.2764 | 1012.6 | 11.2499 | 1023.7 | 11.217 | 1034.8 | 11.1781 | 1045.9 | 11.1335
1057 | 11.0835 | 1068.1 | 11.0284 | 1079.2 | 10.9686 | 1090.3 | 10.9043 | 1101.4 | 10.8358
1112.5 | 10.7635 | 1123.6 | 10.6875 | 1134.7 | 10.6082 | 1145.8 | 10.5257 | 1156.9 | 10.4404
1168 | 10.3524 | 1179.1 | 10.2619 | 1190.2 | 10.1693 | 1201.3 | 10.0746 | 1212.4 | 9.9781
1223.5 | 9.88 | 1234.6 | 9.7804 | 1245.7 | 9.6795 | 1256.8 | 9.5774 | 1267.9 | 9.4744
1279 | 9.3705 | 1290.1 | 9.2659 | 1301.2 | 9.1607 | 1312.3 | 9.055 | 1323.4 | 8.949
1334.5 | 8.8427 | 1345.6 | 8.7363 | 1356.7 | 8.6299 | 1367.8 | 8.5234 | 1378.9 | 8.4172
1390 | 8.3111 | 1401.1 | 8.2053 | 1412.2 | 8.0998 | 1423.3 | 7.9948 | 1434.4 | 7.8903
1445.5 | 7.7862 | 1456.6 | 7.6828 | 1467.7 | 7.58 | 1478.8 | 7.4779 | 1489.9 | 7.3765
1501 | 7.2759 | 1512.1 | 7.176 | 1523.2 | 7.077 | 1534.3 | 6.9789 | 1545.4 | 6.8816
1556.5 | 6.7852 | 1567.6 | 6.6898 | 1578.7 | 6.5954 | 1589.8 | 6.5019 | 1600.9 | 6.4093
1612 | 6.3178 | 1623.1 | 6.2273 | 1634.2 | 6.1378 | 1645.3 | 6.0494 | 1656.4 | 5.962
1667.5 | 5.8756 | 1678.6 | 5.7903 | 1689.7 | 5.706 | 1700.8 | 5.6228 | 1711.9 | 5.5407
1723 | 5.4596 | 1734.1 | 5.3795 | 1745.2 | 5.3005 | 1756.3 | 5.2226 | 1767.4 | 5.1457
1778.5 | 5.0699 | 1789.6 | 4.9951 | 1800.7 | 4.9213 | 1811.8 | 4.8485 | 1822.9 | 4.7768
1834 | 4.7061 | 1845.1 | 4.6364 | 1856.2 | 4.5677 | 1867.3 | 4.5 | 1878.4 | 4.4333
1889.5 | 4.3675 | 1900.6 | 4.3027 | 1911.7 | 4.2389 | 1922.8 | 4.176 | 1933.9 | 4.1141
1945 | 4.053 | 1956.1 | 3.9929 | 1967.2 | 3.9337 | 1978.3 | 3.8754 | 1989.4 | 3.818
2000.5 | 3.7615 | 2011.6 | 3.7058 | 2022.7 | 3.651 | 2033.8 | 3.597 | 2044.9 | 3.5438
2056 | 3.4915 | 2067.1 | 3.44 | 2078.2 | 3.3893 | 2089.3 | 3.3393 | 2100.4 | 3.2902
2111.5 | 3.2418 | 2122.6 | 3.1941 | 2133.7 | 3.1472 | 2144.8 | 3.1011 | 2155.9 | 3.0557
2167 | 3.0109 | 2178.1 | 2.9669 | 2189.2 | 2.9236 | 2200.3 | 2.881 | 2211.4 | 2.839
2222.5 | 2.7977 | 2233.6 | 2.757 | 2244.7 | 2.717 | 2255.8 | 2.6776 | 2266.9 | 2.6389
2278 | 2.6007 | 2289.1 | 2.5632 | 2300.2 | 2.5262 | 2311.3 | 2.4899 | 2322.4 | 2.4541
2333.5 | 2.4188 | 2344.6 | 2.3842 | 2355.7 | 2.35 | 2366.8 | 2.3165 | 2377.9 | 2.2834
2389 | 2.2509 | 2400.1 | 2.2189 | 2411.2 | 2.1874 | 2422.3 | 2.1564 | 2433.4 | 2.1258
2444.5 | 2.0958 | 2455.6 | 2.0662 | 2466.7 | 2.0371 | 2477.8 | 2.0085 | 2488.9 | 1.9803
2500 | 1.9525 | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
#N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
#N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
#N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
#N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
2500 | 1.9525