## Summary
Service Report
COMPANY |  |  | Cal Tech
CUSTOMER |  |  | Bruce Brunschwig
DATE |  |  | 2003-12-11
ENGINEER |  |  | Scott
Work completed
1. | Replaced sample introduction arm flange.
2. | Aligned intro arm to sample transfer stages.
3. | Repaired sample retention clip on chamber sample stage platform.
4. | Installed sample stage motor controller under vacuum chamber tabletop.
5. | Replaced 2503 Memory Controller power supply capacitors.
6. | Replaced 8724 Spectrometer Power Supply capacitors and cooling fan.
7 | Bi-annual instrument service visit.
8 | Installed latest version of ESCA 2000E Capture and Analysis software.
Follow-up Items
1. | Find cause of low ESCA count rate.
2. | Prep chamber gate valve leaks closed and prep chamber is vented.

## 8724
8701 DATA SHEET
COMPANY |  | Cal Tech
CUSTOMER |  | Bruce Brunschwig
DATE |  | 2003-12-11
TEST ENGINEER |  | Scott
PARAMETER |  | PE 25eV | PE 50eV | PE 100eV | PE 150eV
B.E. |  | 84 | 84 | 84 | 84
Vo |  | 1378.7 | 1351.8 | 1296.7 | 1242.7
BOOK |  | 1368.7 | 1340.3 | 1282.9 | 1226
DELTA |  | -10 | -11.5 | -13.8 | -16.7
SPHERE + |  | 1365.9 | 1326.5 | 1246.2 | 1166.9
BOOK |  | 1356.2 | 1315.5 | 1233.6 | 1152.4
ADJ VALUE |  | 1355.9 | 1315 | 1232.4 | 1150.2
SPHERE - |  | 1391.3 | 1376.7 | 1346.9 | 1317.8
BOOK |  | 1381.1 | 1365 | 1332.1 | 1299.4
BOOK SPHERE |  | 24.9 | 49.5 | 98.5 | 147
CALC SPHERE |  | 25.4 | 50.2 | 100.7 | 150.9
I Mag |  | 52 | 72 | 106 | 131
BOOK |  | 53 | 75 | 106 | 130
V trim |  | 1383.9 | 1362.8 | 1318.9 | 1275.9
BOOK |  | 1374.4 | 1352 | 1306.7 | 1261.1
ADJ VALUE |  | 1373.9 | 1351.3 | 1305.1 | 1259.2
V1 |  | 904.9 | 944.9 | 965 | 976.1
BOOK |  | 750 | 825 | 900 | 1000
VQ |  | 1394.2 | 1383.9 | 1362.8 | 1341.8
BOOK |  | 1382.9 | 1370 | 1343.3 | 1316.5
ADJ VALUE |  | 1384.2 | 1372.4 | 1349 | 1325.1
V2 |  | 1442.6 | 1480.3 | 1556.6 | 1631.8
BOOK |  | 1431.7 | 1469.7 | 1543.4 | 1615.3
ADJ VALUE |  | 1432.6 | 1468.8 | 1542.8 | 1615.1
SET BINDING ENERGY TO 1100 Ev
V1 |  | 402 | 301.4 | 326.5 | 337.4
V0 |  | 356.6 | 329.7 | 274.8 | 222.9
SLOPE |  | 0.49203 | 0.62959 | 0.62482 | 0.6263
INTERCEPT |  | 226.5434 | 93.8254 | 154.8004 | 197.7979
V TEST |  | 8.919
V HRES |  | 3.6
V LRES |  | 43 |  |  |  | ANALOG BORD MODEL PARAMETERS
|  | 3.2896 | 6.9535 | 6.7483 | 6.8111 | R (V0-V1) meg | 18.8672 | OS(25eV) | 83.333 | R3 | 100
R2 (in MEGOHMS) |  | 1 | 4 | 4 | 4 | R(V1-Grd) meg | 14.62 | OS(50eV) | 166.67 | R4 | 0.89993
INTERCEPT OUT |  | 504.1486 | 195.4739 | 256.985 | 251.9512 | RA meg | 0.5 | OS(100eV) | 333.33
[6 empty] | RB meg | 0.5 | OS(150eV) | 500
[6 empty] | R1 meg | 3
[6 empty] | R2 INSTALLED | 2
[6 empty] | R ab | 0.25
[6 empty] | R red | 2.1991
[6 empty] | K1 | 1.0833

## 8702
8702 PARAMETERS
COMPANY | Cal Tech
CUSTOMER | Bruce Brunschwig
DATE | 2003-12-11 |  | COMMENT
TEST ENGINEER | Scott
SPOT | CURRENT | PIERCE | F1 | SPOT L | CALC L | SPOT W | PHOTO e
OFF | 2 | 914 | 7682
150 | 3.5 | 420 | 8247 |  | 0
300 | 7.5 | 339 | 8283 |  | 0
600 | 12 | 282 | 8464 |  | 0
800 | 22 | 152 | 8580 |  | 0
FILAMENT/ CATHODE PARAMETERS
| I 2kV | E 2kV | K POWER |  | I FIL |  | E 10kV
| 5.6 | 2.209 | 12.3704 |  | 1.6 |  | 10.02

## Xprobe
X-Probe Performance Specification
COMPANY | Cal Tech
CUSTOMER | Bruce Brunschwig
DATE | 2003-12-11 |  |  | COMMENT | Au 60%
TEST ENGINEER | Scott
Unscanned Au 4f 7/2 One minute acquistion
RESOLUTION | SPOT SIZE
| 150 micron |  | 300 micron |  | 600 micron |  | 800 micron
| SPEC | ACT | SPEC | ACT | SPEC | ACT | SPEC | ACT
RES 1 - 25eV
RESOLUTION | 0.69 | 0.89 | 0.75 |  | 0.95 |  | 1.1 | 0.9
AREA - K CNT | 10 | 6.61 | 55 |  | 100 |  | 140 | 41.07
RES 2 - 50eV
RESOLUTION | 0.85 |  | 0.9 | 0.87 | 1.1 |  | 1.2
AREA - K CNT | 30 |  | 150 | 43.87 | 260 |  | 335
RES 3 - 100eV
RESOLUTION | 1.25 |  | 1.3 |  | 1.4 | 1.21 | 1.5
AREA - K CNT | 50 |  | 250 |  | 410 | 106.16 | 510
RES 4 - 150eV
RESOLUTION | 1.6 | 1.49 | 1.65 |  | 1.7 |  | 1.8 | 1.73
AREA - K CNT | 60 | 14.93 | 280 |  | 470 |  | 660 | 128.4
Scanned Au 4f 7/2 one scan - Area in thousands counts
RESOLUTION | SCAN W | SPOT SIZE | FWHM | FWHM | AREA | AREA | PEAK BE | GAUSS %
| eV | MICRON | SPEC | ACTUAL | TYPICAL | ACTUAL | eV
RES 1 - 25eV | 10 | 150 | 0.77 | 0.93 |  | 414.43
RES 2 - 50eV | 10 | 150 | 0.99 | 1.02 |  | 1661.73
RES 2 - 50eV | 10 | 300 | 1 | 0.9 |  | 6166.41
RES 4 - 150eV | 20 | 150 | 1.9 | 1.77 |  | 7786.73
RES 4 - 150eV | 20 | 800 | 1.9 | 2.06 |  | 57908.5

## Continuity
Analyzer Continuity
COMPANY | Cal Tech
CUSTOMER | Bruce Brunschwig
DATE | 2003-12-11
ENGINEER | Scott
TEST |  | Typical | Measured
Det+ to Det- cables 8701 |  | 15-20 M | 16.96 M
V1 to Ground |  | 12-20 M | 16.18 M
V2 to Ground |  | Open | Open
Vq to Ground |  | Open | Open
V0 to Ground |  | 30-50 M | 32.65 M
Ground to 8701 chassis |  | <0.7 ohm | 0.23 ohsm
V2 to Vq |  | Open | Open
V0 to V1 |  | 15-20 M | 16.47 M
V0 to Vq |  | Open | Open
V0 to Imag |  | 7-10 ohm | 10.05
V0 to Vtrim |  | Open | Open
V+ to V- |  | 3-6 M | 3.52 M
V- to Vtrim |  | 1.2-2.5 M | 1.33 M
V+ to Vtrim |  | 2-4 M | 2.13 M
Vtrim to Det - |  | Open | Open

## Service Checklist
Maintenance Checks
COMPANY |  | Cal Tech
CUSTOMER |  | Bruce Brunschwig
DATE |  | 2003-12-11
ENGINEER |  | Scott
|  | 6 MONTH | Performed | 12 MONTH | Performed | COMMENTS
ELECTRONICS
Clean fan filters |  |  |  | X | X
Clean rack filter |  |  |  | X | X
Check lamps |  | X |  | X | X
Check interlocks
| Vacuum | X |  | X | X
| Air pressure | X |  | X | X
| Transfer | X |  | X
| Anode flow | X |  | X | X
Optimize detector voltage |  |  |  | X | X
Check Signature |  | X |  | X
Calibrate 8702 meter |  |  |  | X | X
Document 8702 |  |  |  | X | X
Document 8701/8724 |  |  |  | X | X
PHYSICS SYSTEM
Replace Anode |  | X |  | X | X
Check Gun Cooling |  | X |  | X | X
Check lens screens |  | X |  | X | X
Test Crystal set point |  |  |  | X
VACUUM SYSTEMS
Record analyzer Base P |  | X |  | X | X |  | 6.8 x 10-9 Torr
Record He pressure (op) |  | X |  | X | X
Record Cryo temp |  | X |  | X | X |  | < 15 K
Replace Absorber |  |  |  | X
Record ion pump current |  | X |  | X
Record Prep base P |  | X |  | X | X |  | 4.49 X 10-9
Change Turbo oil |  |  |  | X | X
Change Mech pump oil |  |  |  | X
Clean exaust oil seperator |  |  |  | X
Clean Fore line filter |  |  |  | X
ANODE COOLING
Replace Ion Exchange filter |  | X |  | X | X
Replace Partical filter |  | X |  | X | X
Top off water |  | X |  | X | X
Set flow to 80 GPH |  | X |  | X | X
Check for water leaks |  | X |  | X | X