Use of Titanium Tetrachloride as a Test for Water

- Pass a bottle of titanium tetrachloride into glove box

- Open the bottle inside the glove box, but be prepared to close the bottle immediately if you see smoke. For best results, hold a black-gloved hand behind the bottle: the white smoke will be better seen against the black back- ground.

- If smoke is observed, moisture exists at a relatively high level .

- If no smoke is observed, the dew point is less than - 6ooc (1o ppm).

Use of Diethylzinc as aTest for Oxygen

Materials used in the following test are pyrophoric when mixed -- use extreme caution in handling.

- Pass a bottle of diethylzinc, a bottle of heptane, and an empty bottle with cap into glove box.

- Mix a small amount of each material together in the empty bottle, but be prepared to cover the bottle immediately if smoke is seen. For best results, hold a gloved hand behind the bottle: the white smoke will be better seen against the black background.

  - If there is less than 5 ppm of oxygen within the glove box, the mixture will not emit smoke.

  - If smoke is seen, a relatively high amount of oxygen exists. Cap bottle immediately.

Use of [TiCp2(NCCH3)2]+  as a Colorimetic Indicator for Oxygen

[TiCp2(NCCH3)2]+  can be used as a colorimetric indicator for oxygen contamination in inert-atmosphere glove boxes. Small quantities of the blue solution are conveniently prepared by dissolving small (mg) amounts of TiCp2Cl2 in acetonitrile in a scintillation vial, adding zinc dust, and capping. Aliquots of this solution may be removed by disposable pipet and transferred to an open vial. Inspection of the aliquot gives a qualitative indication of the dioxygen level in the glove box. As the solution is exposed to trace O2 the solution color changes from blue to green and ultimately, under higher dioxygen concentrations, to yellow.

J. of Chem. Ed, 1998, 75, 459.