Braun

Catalyst: 

1. lasts 5 – 10 years consists of a copper oxide on Alumina catalyst and Molecular Sieves 

2. Need: 4.5 kg of 14x20 particle size (or 8x14) Cu catalyst

a. Braun FC-01 $425

b. Englehard: Cu-0226 <425

3. 12 lbs 13X beads (not pellets) We used 8-12  mesh .  Molecular sieve FC-02$150.00

Accessories

1. Spare Gloves

2. Static Eliminator part ACC-11 $460

3. Internal glove port cover  BP-03

4. PLC H2O analyzer

5. PLC O2 analyzer

Regeneration

1. Purge with Regen gas for 345 mins

a. Rg open

b. E open

2. Evacuate for 60 mins

a. Rg closed

b. E closed

c. Rv open

3. Cool and cycle for 660 minuets

a. Rv closed, K opens

b. Cycles through the opening and closing

Englehard Corp 800 321 2747 120 Pine St POB 4017; Elyria, Ohio

Mbraun
603 773 9333 14 Marin Way Stratham NH

Working gas:  N2(g) from high pressure Liq N2 tank

Regeneration gas:
Argon – H2  or N2 – H2 5 – 10 % H2 at 7.5 psi

Regeneration:

1. Switch off the circulation operation by turning off the blower.  Green light goes out. Blower switches off.

2. Close hand-valves on purifier columns

3. Before starting the regeneration it is recommended to check the program manually:  Set regeneration program knob at “purging”, purging-valve for regeneration gas should open. Set the pressure at the  pressure reduction valve at a value which give a flow rate of 15 -  20 l/min.  This is equivalent to a pressure of about 7 psi.

4. After setting the regenerating gas purging amount, turn the regenerating program dial to the STOP position.

5. Continue to turn the program dial clockwise until the regeneration lights ups (red indicator).  The program runs to  its end automatically.  Do not turn the program dial while the program is in progress, in order to avoid damage to the system.  The end of the e regeneration program is indicated by extinguishing of the indicator REGENERATION.  A complete regeneration cycle takes around 10 hours.  The unit is then ready for operation again.

6. Open the hand valves on the purifier.

7. Using the blower switch put the system back into normal operation.

