###

___abc_________________________________________________________________

Description: Use abc to perform an automatic baseline correction.

The abc command automatically selects data points to

pull to the new baseline.  (To choose your own data

points, use the base command.)

Returns:     1 for success, 0 for failure.  The return value is

also placed in the z register.

Arguments:   None.

Example:     load "co.abs"  /* Load file into memory. */

abc            /* Correct the baseline. */

show           /* Display the spectrum. */

###

___abs2ras_____________________________________________________________

Description: Use to convert absorbance data (the current sample) to

transmittance data.

Arguments:   None.

Returns:     1 for success, 0 for failure.  The return value is

also placed in the z register.

Example:     load "co.abs"  /* Load file into memory. */

abs2ras        /* Convert data to transmittance. */

show           /* Display the transmittance data. */

###

___acos________________________________________________________________

Description: Use to compute the arc cosine of the specified value.

Arguments:   Value (in radians) for which to find arc cosine (int or

float).

Returns:     Result (float).  The return value is also placed in the

z register.

Example:     y = acos 0   /* Place arc cosine of 0 in y */

/* register. */

x = acos z   /* Place arc cosine of value */

/* in register z into register x. */

Also See:    asin, atan, cos, sin, tan

###

___add_________________________________________________________________

Description: Use to add data to the data in the memory.  The

resolution of the data you add must be the same

as the resolution of the data in memory.

Arguments:   Name of file you wish to add (string).

Returns:     1 for success, 0 for failure.  The return value is

also placed in the z register.

Example:     load "pure1.abs"       /* Load file into memory. */

add "pure2.abs"        /* Add pure2 to pure1. */

title "pure1 + pure2"  /* Assign title to result. */

save "puresum.abs"     /* Save the result. */

show                   /* Display the data. */

Also See:    average, divide, multiply, subtract

###

___add2Lib_____________________________________________________________

Description: Add a data file to a user library.

Arguments:   Name of library (string), name of data file to

add (string).

NOTE: If you do not include a data file name, the

current sample is added.

Returns:     1 for success, 0 for failure (int).  (A random value is

placed in the z register.)

Example:     load "pure1.abs"   /* Load file into memory. */

add2lib "purelib"  /* Add sample to library. */

add2lib "purelib" "pure2.abs"

/* Add pure2.abs to library. */

Also See:    library

###

___advancedCall________________________________________________________

Description: Use from within a macro to place an advanced call like

those you can define for and use from the WinFIRST

Advanced menu.

NOTE: The WinFIRST Advanced menu and its use is

described in detail in "Chapter 12  The WinFIRST

Advanced Menu" in your "WinFIRST User's Reference

Guide."  Please look there for further information.

Arguments:   The program name, the command line template, the

program type.  (See Chapter 12 in your WinFIRST

User's Guide for further information.)

Returns:     1 for success, 0 for failure (int).

Example:     a = advancedCall "notepad.exe" "%t" 2

###

___annotate____________________________________________________________

Description: Annotate is an object that contains properties and

methods you can set and execute to perform various

annotation operations.

Properties:  Properties for the annotate object are:

off:           Use to make existing annotations invisible on the

data display.

Arguments:     None.

Syntax:        annotate:off

on:            Use to make existing annotations visible on the

data display.

Arguments:     None.

Syntax:        annotate:on

textwidth:     Use to set or get the text width displayed when

using the Edit Annotations dialog.

Arguments:     The maximum number of characters to display.

Maximum value allowed is 127.

Set Syntax:    annotate:textwidth=32

Get Syntax:    i=annotate:textwidth

Methods:     Methods for the annotate object are:

clear:         Use to remove a specified annotation.  (The

annotation is specified by a number that

represents the annotation location in the list

displayed using the list method.)

Arguments:     The number of the annotation to remove (int).  If

no value is specified, or if 999 or 0 is

specified, all annotations are removed.

Syntax:        annotate:clear 1  /* Clear first annotation. */

list:          Use to list the annotations for the current

sample.

Arguments:     None.

Syntax:        annotate:list

Also See:    drawbox, drawHruler, drawline, drawtext, drawVruler

###

___append______________________________________________________________

Description: Use to move to the end of a text or binary file that

is currently open for input and output.

NOTE: To avoid overwriting text or data in an open

file, use append before writing to that file.

Arguments:   Number assigned to open file (int).

Returns:     1 for success, 0 for failure (int).

Example:     fopen 1 "file1"  /* Open file1 for read and write. */

append 1         /* Move to the end of file1. */

fprint 1 "Place this at end of file1"

/* Append text to the file end. */

fclose 1         /* Close the file. */

Also See:    eof, exist, fclose, filepos, finput, fopen, fprint,

fread, freadByte, fseek, fwrite, fwriteByte, rewind

###

___ArrayToData_________________________________________________________

Description: Use to find the wavenumber value for a given array

position.

Arguments:   Array position (int), array to use (variable).

Returns:     Wavenumber associated with position (float).  The

return value is also placed in the z register.

Example:     load co.abs  /* Load file into memory. */

x = arraytodata 10 sample:data

/* Get wavenumber at position 10 */

/* and place value in register x. */

print x      /* Print the value to the screen. */

Also See:    dataToArray

###

___asin________________________________________________________________

Description: Use to compute the arc sine of the specified value.

Arguments:   Value (in radians) for which you want the sine (int

or float).

Returns:     Result (float).  The return value is also placed in

the z register.

Example:     y = asin 0   /* Place arc sine of 0 in y */

/* register. */

x = asin z   /* Place arc sine of value stored */

/* in z register into x register. */

Also See:    acos, atan, cos, sin, tan

###

___ask_________________________________________________________________

Description: Use to display a yes or no question and wait for a

response.

Arguments:   (optional)  Text of question to display (string).

Returns:     1 if y is typed, 0 if y is not typed (int).  Also, the

return value is stored in the z register.

Example:     float array tmpdata  /* Declare a float array. */

load "co.abs"        /* Load file into memory. */

ask "This macro alters the sample data... Continue?"

if (z == 0)  /* If answer is not y... */

exit       /* Exit the macro. Otherwise, continue. */

tmpdata:data = sample:data

/* Assign sample data to the array */

/* named tmpdata. */

tmpdata:firstx = 400.0  /* Set starting wavenumber. */

tmpdata:lastx = 4000.0  /* Set ending wavenumber. */

sample:data = tmpdata:data

/* Assign tmpdata to be the sample. */

autoscale    /* Display sample with new range. */

###

___atan________________________________________________________________

Description: Use to compute the arc tangent of the specified value.

Arguments:   The value (in radians) for which to calculate the arc

tangent (int or float).

Returns:     Result (float).  The return value is also placed in

the z register.

Example:     y = atan 0   /* Place arc tangent of 0 in */

/* register y. */

x = atan z   /* Place arc tangent of value in */

/* register z into register x. */

Also See:    acos, asin, cos, sin, tan

###

___autoScale___________________________________________________________

Description: Use to rescale sample data to the data limits.

(Autoscale does not undo any data manipulations.)

Arguments:   None.

Returns:     0 (int).  The return value is also placed in the z

register.

Example:     load "pure1.abs"          /* Load file into memory. */

expand 3000 3500 0.5 1.5  /* Expand the display. */

show                      /* Show the results. */

autoscale                 /* Rescale data display. */

Also See:    autoscaleX, autoscaleY, display:expandXOnly,

display:rescaleToLimits, expand, expandX, expandY

###

___autoScaleX___________________________________________________________

Description: Use to rescale sample data to the data limits, while

preserving the current y axis display limits.

Arguments:   None.

Returns:     1 for success, 0 for failure.  The return value is

also is placed in the z register.

Example:     load "pure1.abs"          /* Load file into memory. */

expand 3000 3500 0.5 1.5  /* Expand the display. */

show                      /* Show the results. */

autoscalex                /* Rescale x axis of data */

/* display while leaving */

/* y axis unchanged.)

Also See:    autoscale, autoscaleY, display:expandXOnly,

display:rescaleToLimits, expand, expandX, expandY

###

___autoScaleY___________________________________________________________

Description: Use to rescale sample data to the data limits, while

preserving the current x axis display limits.

Arguments:   None.

Returns:     1 for success, 0 for failure.  The return value is

also is placed in the z register.

Example:     load "pure1.abs"          /* Load file into memory. */

expand 3000 3500 0.5 1.5  /* Expand the display. */

show                      /* Show the results. */

autoscaley                /* Rescale y axis of data */

/* display while leaving */

/* x axis unchanged.)

Also See:    autoscale, autoscaleX, display:expandXOnly,

display:rescaleToLimits, expand, expandX, expandY

###

___average_____________________________________________________________

Description: Use to average a data file you specify with the sample

data loaded in memory.

Arguments:   Name of file to average (string) with sample data.

Returns:     0 (int).  The return value is also placed in the z

register.

Example:     load "pure.abs"      /* Load file into memory. */

average "pure2.abs"  /* Average data in pure.abs */

/* and pure2.abs files. */

show                 /* Display the result. */

Also See:    add, divide, multiply, subtract

###

___background__________________________________________________________

Description: Background is an object that contains methods you can

set and execute to load, save and show a current

background data file.

Methods:     Methods for the background object are:

load:          Use to load a background data file to serve as the

current background data file.  The loaded file

must be one of these data types: .igm, .big, .sbm

or .bkg.

Arguments:     The name of the background data file to load

(string).

Syntax:        background:load pure.bkg

save:          Use to save the current background data file.

Arguments:     The name of the background data file to save

(string).

Syntax:        background:save pure2.bkg

show:          Use to display the current background data file

in a data window.

Arguments:     None.

Syntax:        background:show

###

___base________________________________________________________________

Description: Use to correct the baseline at specified wavenumbers.

Arguments:   At least 2, and up to 19, wavenumbers (float).

Returns:     0 (int).  The return value is also placed in the z

register.

Example:     load "co.abs"        /* Load file into memory. */

base 2250 2100 2050  /* Correct baseline at these */

/* wavenumbers. */

show                 /* Display the result. */

###

___basename____________________________________________________________

Description: Use to remove the filename extension from a file name.

Arguments:   File name, including filename extension (string).

Returns:     The filename, with no filename extension (string).

Example:     load "pure.abs"          /* Load file into memory. */

x = basename "pure.abs"  /* Assign filename, or */

/* basename, to register x. */

title x            /* Set title of current sample to */

/* the string in register x. */

###

___beep________________________________________________________________

Description: Use to make the computer to beep.

Arguments:   None.

Returns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:     beep     /* Beep once. */

sleep 1  /* Sleep for 1 second. */

beep     /* Beep a second time. */

ask "Do you want to initialize the bench?"

/* Ask question and wait for answer. */

###

___bench_______________________________________________________________

Description: Bench is an object that contains properties and methods

that you can set and execute to perform various

spectrometer functions.

Properties:  You can get the current setting for any property

specified by bench.  For example:

r = bench:method_name  /* Store current setting */

/* for the method_name */

/* property in r register. */

You can also set a new setting for any property

specified by bench.  For example:

bench:method_name = "scnmeth1"

/* Set setting for method_name */

/* property to scnmeth1. */

Properties for the bench object are:

background_scans: Use to specify the number of background scans

to collect.

Arguments:     The number of scans (int).

Syntax:        bench:background_scans = 20

baud_rate:     This is a read-only property.  Use it to report

the current baud rate setting for computer-

spectrometer communication, as specified by the

current data collection method.

Syntax:        a = bench:baud_rate;print a

benchfft:      Use to specify where the FFT is performed.

Arguments:     1 for FFT on spectrometer, 0 for FFT on computer

(int).

Syntax:        bench:benchfft = 1

communications_port: This is a read-only property.  Use it to

report the current communications port setting to

use for data collection, as specified by the

current data collection method, where:

1 = COM1, 2 = COM2, 3 = COM3 and 4 = COM4.

Syntax:        a = bench:communications_port;print a

data_directory_mode: Use to specify the data directory to use.

NOTE: You can use the cwd property of the parms

object to report the name of the current data

directory, that is, the current working directory.

Arguments:     The number that represents the directory to use

(int), where:

0 = the current month directory.

1 = a directory with the same name as the current

data collection method.  If the directory does not

exist it is created.

2 = the directory specified by the data collection

method.  If the directory does not exist it is

created.

Syntax:        bench:data_directory_mode = 0

data_type:     Use to specify the type of data you want to

collect.

Arguments:     The number that represents the data type (int),

where: 0 = abs, 1 = ras, 2 = drt, 3 = sbm, 4 =

bkg, 5 = igm and 6 = big.

Syntax:        bench:data_type = 0

ending_frequency: Use to specify the data's ending wavenumber.

Arguments:     A wavenumber (float).

Syntax:        bench:ending_frequency = 400

file_naming:   Use to specify the file naming mode to use.

Arguments:     A number that represents the naming mode (int),

where:

0 = None.  (Scan file name is used without

increment value.)

1 = Prompt for entry of name (during collection).

2 = Auto increment.  (The name specified with the

scan_file_name property and an increment are used

to assemble filename.)

4 = Default.  (An eight-character filename is

assembled using month, day and time of day.  For

example, 01081630.abs was created on Jan. 8 at

4:30 p.m.)

Syntax:        bench:file_naming = 0

fvelindex:     Use to specify the forward scan velocity.

NOTE: For Genesis I FTIR spectrometers, use only

index 2 (6.25 KHz) or index 4 (25 KHz).

WARNING: When using WinFIRST with Mattson

spectrometers other than Genesis I, refer to

settings in the software for forward scan

velocity.  (Use the Advanced page of the control

panel's Method Setup dialog.)

Arguments:     The velocity index to use.

Syntax:        bench:fvelindex = 2

increment:     Use to specify the number to use as the starting

increment when Auto increment file naming is used.

Arguments:     The starting increment value (int).

Syntax:        bench:increment = 5

lpfindex:      Use to select the internal or external lowpass

filter.

NOTE: For Genesis I FTIR spectrometers, for a scan

velocity of 6.25 KHz, use the internal filter.

For a scan velocity of 25 KHz (and with an MCT

detector in use), use an external filter.

Arguments:     The filter index to use (int), where: 0 = internal

and 1 = external.

Syntax:        bench:lpfindex = 0

method_name:   Use to specify the name of the scan method to use.

(The scan method is also called the data

collection method.)

Arguments:     The scan method filename (string).

Syntax:        bench:method_name = special

prompt_for_title: Use to specify whether or not you are prompted

to enter a sample data file title.

Arguments:     1 to enable or 0 to disable (int).

Syntax:        bench:prompt_for_tiltle = 0

realtime_coadd_mode: Use to specify whether co-add stops at a

specified number of scans or at a specified

signal-to-noise threshold.

Arguments:     0 for number of scans, 1 for snr threshold (int).

Syntax:        bench:realtime_coadd_mode = 0

reference_time_threshold: Use to specify the length of time, in

minutes, beyond which the background spectrum is

considered invalid.

Arguments:     A number to indicate threshold in minutes (int).

Syntax:        bench:reference_time_threshold = 3

resolution:    Use to specify the resolution of the spectral data

to collect.

Arguments:     The wavenumber resolution (float).  Valid values

are: 1.0, 2.0, 4.0, 8.0, 16.0, 32.0, 64.0, and

128.0 wavenumbers.

Syntax:        bench:resolution = 4.0

rvelindex:     Use to specify the reverse scan velocity.

NOTE: For Genesis I FTIR spectrometers, use only

index 2 (6.25 KHz) or index 4 (25 KHz).

WARNING: When using WinFIRST with Mattson

spectrometers other than Genesis I, refer to

settings in the software for reverse scan

velocity.  (Use the Advanced page of the control

panel's Method Setup dialog.)

Arguments:     The velocity index to use.

Syntax:        bench:rvelindex = 2

sample_scans:  Use to specify the number of sample scans to

collect.

Arguments:     The number of scans (int).

Syntax:        bench:sample_scans = 50

saveRawData:   Use to specify whether raw interferogram data

is saved.

Arguments:     1 to save raw data, 0 to not save raw data (int).

Syntax:        bench:saverawdata = 0

scan_file_name: Use to specify the base name to use with auto-

increment file naming.

Arguments:     The name to use (string).

Syntax:        bench:scan_file_name = tmp

scan_mode:     Use to specify the data collection mode to use.

Arguments:     The number that represents the scan mode (int),

where: 0 = Standard FTIR, 1 = COADD FTIR, 2 =

Standard Microscope, 3 = COADD Microscope.

Syntax:        bench:scan_mode = 0

scanPromptLevel: Use to specify the amount of prompting to show

during data collection.

Arguments:     The number that represents the scan prompt level

(int), where: 0 = low and 1 = high.

Syntax:        bench:scanpromptlevel = 0

scanVerbosityLevel: Use to specify the amount of informational

messages that WinFIRST.

Arguments:     The number that represents the scan verbosity

level (int), where: 0 = silent (no messages),

1 = brief (some messages), and 2 = normal (all

messages).

Syntax:        bench:scanverbositylevel = 0

signal_threshold: Use to specify the signal-to-noise threshold

to use (if the realtime_coadd_mode property is set

with 1 as its argument).

Arguments:     A number to serve as the threshold (int).

Syntax:        bench:signal_threshold = 5

starting_frequency: Use to specify the starting wavenumber for

the data to collect.

Arguments:     A wavenumber (float).

Syntax:        bench:starting_frequency = 500

test:          Use to run diagnostic communication and

spectrometer tests up to and including the test

level specified in the scan method.

Arguments:     None.

Syntax:        bench:test

tune_points:   Use to specify the total number of data points to

collect as a tune interferogram.

Arguments:     The number of data points (int).

Syntax:        bench:tune_points = 200

useShuttle:    Use to specify whether a sample shuttle is in use.

Arguments:     1 to use shuttle, 0 to not use shuttle (int).

Syntax:        bench:useshuttle = 0

writeBkg:      Use to specify whether the collected background is

saved.

Arguments:     1 to save background, 0 to not save background

(int).

Syntax:        bench:writebkg = 1

writeSample:   Use to specify whether the collected sample is

saved.

Arguments:     1 to save sample, 0 to not save sample (int).

Syntax:        bench:writesample = 1

Methods:     Methods for the bench object are:

autoAlign:     Use to perform a noninteractive auto-align (tune)

procedure.

Arguments:     None.

Syntax:        bench:autoalign

bkg:           Use to collect a background (using the current--

most recently loaded-�data collection method).

Arguments:     None.

Syntax:        bench:bkg

findAcc:       Use to report the position of a spectrometer

accessory you identify as an argument.

NOTE: This method for the bench object works only

with spectrometer firmware for the Galaxy and

later generation FTIR spectrometers.  Consult your

Mattson sales representative for further

information.

NOTE: Also see the moveAcc method for the bench

object.

Arguments:     From the table below, a device index to identify

the spectrometer accessory for which you want to

report position (or existence).

Supported Device Indexes and Positions

Device

Name               Index    Position

Iris                 0      0 - 100 (percent open)

Shuttle              2      1 = sample

2 = background

Calibration Access.  3      1 = sample

(calibration sample in

beam)

2 = normal (sample out

of beam)

X-Axis               4      any (generic stepping

device)

Y-Axis               5      any (generic stepping

device)

Sample Wheel         6      any (of 6 or of 24

positions).  0 is home

position

DRT Wheel            7      (depends on number of

positions on wheel).

0 is home position

Right-Port Mirror    9      1 = external

2 = internal

Left-Port Mirror     12     1 = external

2 = internal

NOTE: Do not use any indexes other than those

listed here.  Some indices are reserved for

internal use, and some others control devices that

are not under user control.  Using device indexes

other than those listed here can have unexpected

and possibly undesirable effects.

Returns:       The position of the specified accessory (see table

above), or -1 if the device is not found.

Syntax:        bench:findAcc 0

Example:       if ((bench:findAcc0) != -1)

/*If an iris is found... */

bench:moveAcc 0 25     /* move it to 25% open. */

findZPD:       Use to instruct the spectrometer to perform a

centerburst scan, in which the ZPD (zero path

difference point) is located.

NOTE: This command is mainly of use for testing

purposes.

Arguments:     None.

Syntax:        bench:findzpd

getRemoteData: Use to retrieve data held in spectrometer

memory.

Arguments:     (optional)  An index value to specify how many

spectra to recieve, where 0 indicates one,

1 indicates two, etc.  An index of 0 indicates the

newest spectrum while 1 indicates the next newest.

The maximum index value varies and depends on the

amount of memory in the spectrometer, the

resolution of spectra, and the number of spectra

held in memory.  When you provide no index, only

the newest spectrum is retrieved.

Syntax:        bench:getremotedata 2

Returns:       1 for success, 0 for failure (int).

Example:       a = 0;

b = 1;

while (b)

{

b = bench:getremotedata a

a = a + 1

}

homeAcc:       Use to move the specifed accessory to the home

position (position 1).

NOTE: Not all accessories support the homeAcc

method.  (See the Arguments section for a listing

of accessories that support the homeAcc method.)

Arguments:     An index to indicate the accessory to move to the

home position, where 4 = the XY stage X axis, 5 =

the XY stage Y axis, 6 = the sample wheel, and 7 =

the DRT wheel.

Syntax:        bench:homeAcc 4

init:          Use to initialize the spectrometer.

Arguments:     None.

Syntax:        bench:init

kscan:         Use to perform a slow kinetics scan (data

collection).

NOTE: When using the bench object with its kscan

method, it is highly recommended that you set up

your system as follows:

1) Under global options, set new data options to

the Replace Sample setting.  So much data is

collected so quickly that Windows can run out of

resources to display new data windows.  (When this

occurs, the data collection continues, but new

data windows are no longer displayed.  In one

test, this occurred after about 150 windows were

displayed.)

2) In the data collection method (choose Method

Setup from the control panel), specify on the data

handling page that sample write (save) is to be

used (check mark), that increments are to be used

to name collected data files (check mark), that

date and time (timestamp) are to be included in

the data file title (under Title Options), and,

optionally, for increased speed, set sample

display to off (no check mark).

Arguments:     (optional)  The number of sample files to collect.

When a value is not specified, the data collection

continues until you press the Escape key to stop

the run.  Even when a number of scans is

specified, you can press the Escape key to stop

the kinetics acquisition.

Syntax:        bench:kscan = 100

moveAcc:       Use to identify a spectrometer accessory and to

specify a position for that accessory.

NOTE: This method for the bench object works only

with spectrometer firmware for the Galaxy and

later generation FTIR spectrometers.  Consult your

Mattson sales representative for further

information.

NOTE: Also see the findAcc method for the bench

object.

Arguments:     From the table below, 1) a device index to

identify the spectrometer accessory and 2) a

position value for that accessory.

Supported Device Indexes and Positions

Device

Name               Index    Position

Iris                 0      0 - 100 (percent open)

Shuttle              2      1 = sample

2 = background

Calibration Access.  3      1 = sample

(calibration sample in

beam)

2 = normal (sample out

of beam)

X-Axis               4      any (generic stepping

device)

Y-Axis               5      any (generic stepping

device)

Sample Wheel         6      any (of 6 or of 24

positions).  0 is home

position

DRT Wheel            7      (depends on number of

positions on wheel).

0 is home position

Right-Port Mirror    9      1 = external

2 = internal

Left-Port Mirror     12     1 = external

2 = internal

NOTE: Do not use any indexes other than those

listed here.  Some indices are reserved for

internal use, and some others control devices that

are not under user control.  Using device indexes

other than those listed here can have unexpected

and possibly undesirable effects.

Returns:       1 for success, 0 for failure (int).

Syntax:        bench:moveAcc 0 25

Example 1:     if ((bench:findAcc0) != -1)

/* If an iris is found... */

bench:moveAcc 0 25  /* move it to 25% open. */

Example 2:     bench:moveAcc 2 2   /* Move shuttle to */

/* background position. */

bkg

bench:moveAcc 2 1   /* Move shuttle to sample */

/* position. */

scan

Example 3:     bench:moveAcc 9 2   /* Move mirror to Internal */

/* position. */

bench:moveAcc 9 1   /* Move mirror to External */

/* position. */

quickCheck:    User to perform a quick "are you there" check

without running through all baud rates (as the

bench:ready method does).  Returns a 1 if

the bench is ready (found) or 0 if not.

Arguments:     None.

Syntax:        bench:quickcheck

readMethod:    Use to load the data-collection method specified

by the method_name property and put in place all

parameters it specifies, including all bench

object properties.

Arguments:     None.

Syntax:        bench:readmethod

ready:         Use to run communication tests, including a full

baud rate test, and return an integer that

identifies the spectrometer type, as follows:

7 = revision 7 spectrometer, 8 = revision 8

spectrometer, 9 = Galaxy or Research Series

spectrometer, 11 = Genesis I spectrometer, 12 =

Infinity spectrometer, 13 = Genesis II

spectrometer and 0 = communication test failure.

Arguments:     None.

Syntax:        bench:ready

saveMethod:    Use to save the current data-collection method,

including all currently specified bench object

properties.

Arguments:     None.

Syntax:        bench:savemethod

scan:          Use to collect a sample (using the current--

most recently loaded--data collection method).

Arguments:     None.

Syntax:        bench:scan

sequence:      Use to perform a multi-range data collection.

Arguments:     (optional)  The name of the sequence configuration

to use.  If you provide no sequence configuration

name, WinFIRST uses the last-used multi-range

sequence configuration.

Syntax:       bench:sequence "example"

setDefaultMethod: Use to tell the spectrometer to save the

current data collection method (current scanning

parameters) as the spectrometer defaults.  (The

spectrometer defaults are the settings that are

put in place each time after you reset the

spectrometer.)  The bench object's

setDefaultMethod method is the only available

way to change the data collection defaults

specified by spectrometer firmware.)

NOTE:The spectrometer does not save accessory

positions when you use this command to save

spectrometer defaults.  At the completion of a

spectrometer reset, accessories return to their

last positions.  Beam selection mirrors return to

the positions that correspond to the selected

detector.

Arguments:     None.

Syntax:        bench:setdefaultmethod

showobj:       Use to report all current settings for all bench

object properties.

Arguments:     None.

Syntax:        bench:showobj

###

___bkg_________________________________________________________________

Description: Use to collect background data.

Arguments:   None.

Returns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:     ask "Would you like to collect a background?"

if (z == 1)  /* If answer is y... */

bkg        /* Collect background data. */

###

___boxcar______________________________________________________________

Description: Use to smooth data using the boxcar method, a straight

running average.

Arguments:   Any odd number between 3 and 25, inclusive, to specify

the number of points to average (int).

Returns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:     load "co.abs"  /* Load file into memory. */

boxcar 9       /* Perform boxcar smooth. */

show           /* Display the results. */

Also See:    boxcar, smooth, smth

###

___break_____________________________________________________________

Description: Use to immediately halt execution of a loop.

Example:     A = 0

while(A < 10)   /* Allow up to 10 executions. */

begin

if (ask "load another reference?")

begin

loadref      /* Load a reference. */

A = A + 1    /* Increment the value of A. */

end

else           /* If answer was 'no' then */

break;       /* stop execution of the loop. */

end

###

___buildString_________________________________________________________

Description: Use to create a character string from any set of

values, including values in variables and registers.

For example, use buildString to create text for output,

text for the ask command, or a string for the

execstring command.

Arguments:   Any list of up to 19 legal literal values, variables,

or registers.

Returns:     All values concatenated in order (string).

Example:     x = "pure"   /* Assign string pure to x. */

ask "Do you want to load the transmittance file?"

/* Ask for type of file. */

if (z == 1)  /* If y is typed...*/

y = ".ras" /* Assign .ras as the extension. */

else         /* Otherwise... */

y = ".abs" /* Assign .abs as the extension. */

f = buildstring x y  /* Build a text string that */

/* contains the filename */

/* stored in x and the filename */

/* extension stored in y. */

load f        /* Load the file with the name that */

/* was created. */

Also See:    execString, getStringElement, isNumber,

setStringElement, string2Float, string2Int,

stringCompare, stringIndex, stringLeft, stringLength,

stringLower, stringMid, stringRight, stringUpper

###

___calibrate___________________________________________________________

Description: Use to analyze the data in the specified file for

slope, intercept, standard deviation and correlation.

Use this command along with the predictX and predictY

commands to implement routines for quantitative

analysis of single components.

NOTE: You must always use the calibrate command before

the predict command.

Arguments:   Name of file containing xy pairs (string).

NOTE: This file must be in ASCII format with one xy

pair per line.

Returns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:     The data file for this example (pairs.dat) contains the

following pairs:

1.0  10.0

2.0  22.0

3.0  31.0

4.0  40.0

int done  /* Declare an integer variable. */

done = 0  /* Set the value of done to 0. */

clrscr    /* Clear the display window. */

print "This program predicts concentration based"

print "on the absorbance amount you enter."

/* Display these messages. */

sleep 5   /* Sleep for five seconds. */

clrscr    /* Clear the screen. */

calibrate "pairs.dat"  /* Calibrate using pairs.dat. */

while (!done)  /* Loop until done is not equal to 0. */

begin     /* Mark top of loop. */

enter a "Enter the absorbance and press Return."

/* Get absorbance value and store in a. */

y = predictY a  /* Predict y using value in a */

/* and store the result in y. */

print "Predicted concentration is at " y

print "when absorbance is " a

/* Print results to screen. */

ask "Do you want to make another prediction?"

if (z == 0)  /* If n is typed... */

done = 1   /* Set done equal to 1. */

end          /* Mark end of loop. */

Also See:    predixtX, predictY, quant

###

___cascadeWindows______________________________________________________

Description: Use to arrange all open data and report windows in a

cascade.

Arguments:   None.

Returns:     1 for success, 0 for failure (int).

Example:     load pure.abs     /* Load data file. */

peaks             /* Generate peak report. */

load pure2.abs    /* Load data file. */

peaks             /* Generate peak report. */

load pstyrene.abs /* Load data file. */

peaks             /* Generate peak report. */

cascadewindows    /* Arrange all open data and */

/* report windows in a cascade. */

Also See:    closeAllWindows, tileWindows

###

___cd__________________________________________________________________

Description: Use to change your current directory.  Unlike the DOS

CD command, this ApPro cd command also lets you

specify a drive as part of the directory designation.

Arguments:   Drive and directory path.

Returns:     The new current directory (string).  Also, 0 is placed

in the z register (int).

Example:     cd \winfirst\data\galaxy  /* Change to the */

/* identified directory. */

Also See:    copy, dir, erase, md, rename

###

___choice______________________________________________________________

Description: In general, use to ask the user a yes-or-no question

and to then continue according to the resultant answer

to the question.  If the user answers "yes," the block

following the "if" statement is executed, otherwise it

is not.

Example:     if(ask "would you like to load a sample")

begin               /* Start of "if" code block. */

load              /* If yes, load a sample. */

end                 /* End of "if" code block. */

/* Otherwise, do nothing. */

Also See:    break, choice and branch, decision loop, for, if, while

Refer also to the expanded discussion of if statements

and other control statements in Chapter 2 of "Using

WinFIRST Macros and ApPro."

###

___choice and branch___________________________________________________

Description: In general, use to ask the user a yes-or-no question

and to then continue according to the results.  If the

user answers "yes," the block following the "if"

statement is executed, otherwise the block following

the "else" statement is executed.

Example:     if(ask "would you like to load a sample")

begin          /* Start of "if" code block. */

load         /* If yes, load a sample. */

end            /* End of "if" code block. */

else

begin          /* Start of "else" code block. */

print "nothing loaded"  /* If no, print message. */

end            /* End of "else" code block. */

Also See:    break, choice, decision loop, for, if, while

Refer also to the expanded discussion of if-else

statements and other control statements in Chapter 2 of

"Using WinFIRST Macros and ApPro."

###

___clearSymbol_________________________________________________________

Description: Use to manually clear the ApPro symbol table.  As

described in Chapter 2 of your ApPro manual, ApPro uses

a symbol table to store user-defined variables and

arrays.  This symbol table is automatically cleared

each time an instruction (one instruction typed at the

ApPro command line or the whole instruction included in

a macro file) is completed.

Arguments:   None.

Returns:     None.

Example:     clearsymbol

###

___clock_______________________________________________________________

Description: Use to report the number of seconds that have elapsed

on the system clock since 00:00:00 Greenwich Mean

Time (GMT), January 1, 1970.  Use the clock command in

combination with the filetime command to determine when

a file was created.  To set a timer to keep track of

seconds, see the timercheck and timerstart commands.

Arguments:   None.

Returns:     The number of seconds.  The return value is also

placed in the z register.

Example:     /* This program determines the age of a file.  Then, */

/* if the file is more than a week old, the program */

/* calls a user-defined function to acquire another */

/* sample: */

int age      /* Declare integer for the file age. */

int current  /* Declare integer for the current time. */

current = clock  /* Get current time and store */

/* value in current. */

age = filetime "tmp.abs"

/* Get age of tmp.abs file. */

if (( current - age ) > 648000)  /* If too old... */

getnewsamp /* Call this user-defined operation. */

/* If file is less than 1 week old (one */

/* week is equal to 604800 seconds), do */

/* nothing. */

###

___closeAllWindows_____________________________________________________

Description: Use to close all open data and report windows.  If

global options specify that WinFIRST is to prompt

you before discarding changed data, then the

save-changed-data prompt appears.

Arguments:   None.

Returns:     1 for success, 0 for failure (int).

Example:     load pure.abs     /* Load data file. */

peaks             /* Generate peak report. */

load pure2.abs    /* Load data file. */

peaks             /* Generate peak report. */

load pstyrene.abs /* Load data file. */

peaks             /* Generate peak report. */

closeallwindows   /* Close all open data and */

/* report windows. */

Also See:    cascadeWindows, tileWindows

###

___clrscr______________________________________________________________

Description: Use to clear the text window.

Arguments:   None.

Returns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:     clrscr  /* Clears all text in the text window. */

###

___cmdargs_____________________________________________________________

Description: Cmdargs is an object that contains the read only

properties of command line arguments passed to a

macro using the run command.

Properties:  Properties for the cmdargs object are:

args:          Use to access an array containing the actual

values passed as arguments to a the macro.

Arguments:     The passed argument's index in the array (int).

Get Syntax:    For all arguments except arrays:

r = cmdargs:args[index]

Returns:       The literal value that was passed.  (The return

type is: cmdargs:type[index].)

Get Syntax:    For arrays:

userarray:data = cmdargs:args[index]

Result:        The userarray takes on all properties of the array

structure passed.

count:         Use to get the number of arguments passed to the current

macro.

Get Syntax:    r = cmdargs:count

Returns:       (int).

type:          Use to access an array that contains the types of

values (integer, string, float, etc.) passed as

arguments to a macro.

Arguments:     The passed argument's index in the array (int).

Get Syntax:    r = cmdargs:type[index]

where: "index" is a number between 1 and

cmdargs:count, inclusive.

Returns:       (string).  Valid return strings are: string, int,

float, byte, char, array and unknown.  (The return

value is "unknown" if an error occurred in passing

arguments.

###

___com1___com2___com3___com4___________________________________________

Description: Four objects, com1, com2, com3 and com4, contain

properties and methods you can use to monitor and

control serial communications.  All four objects have

all of the same properties and methods, as described

below.  Error codes associated with these object are

also described below.

NOTE: If any one of the serial communication ports

(identified as com1, com2, com3 or com4) is open

(after using the corresponding object with the method

named "open"), that port is then unavailable for data

collection.  (The bench object and its property named

"communication_port" identify the port that WinFIRST

software will use, or try to use, for data collection.)

Properties:  NOTE: All properties, with the exception only of

verbose and errCount, should be set when the object

(the serial communication port) is closed.  If the port

is open, the return value is an error code

(-2, IE_OPEN) and the errCount property is not

incremented.

NOTE: Default property values are the settings put in

place at port initialization.

Properties for the com1, com2, com3 and com4 objects

are:

baud:          Use to set or get the baud rate for the object.

Arguments:     300, 600, 1200, 2400, 4800, 9600, 19200, 38400, or

57600.

Set Syntax:    r = com1:baud = 9600

Returns:       If successful, r = 1; if not, r = error code

(int).

Get Syntax:    r = com1:baud

Returns:       The current baud rate for com1.

Default:       38400 (int).

dataBits:      Use to set or get the number of bits per

character.

Arguments:     A number from 4 to 8 (int).

Set Syntax:    r = com1:dataBits = 8

Returns:       If successful, r = 1; if not, r = error code

(int).

Get Syntax:    r = com1:dataBits

Returns:       The current number of data bits for com1 (int).

Default:       8 (int).

errCount:      Use to retrieve the accumulated error count.

Arguments:     0 (resets errCount and lastError).

Set Syntax:    com1:errCount = 0

Returns:       Always sets r = 1 (int).

Get Syntax:    r = com1:errCount

Returns:       The number of errors on com1 since errCount was

reset (int).

flow:          Use to set or get the flow control scheme.

Arguments:     A number to indicate flow control (int), where

0 = none, 1 = CTS/RTS, 2 = DSR/DTR and

3 = XON/XOFF.

Set Syntax:    r = com1:flow = 1

Returns:       If successful, r = 1; if not, r = error code

(int).

Get Syntax:    r = com1:flow

Returns:       The current flow control scheme for com1 (int).

Default:       1 (CTS/RTS) (int).

lastError:     Use to get the last error encountered.

Arguments:     None.

Get Syntax:    r = com1:lastError

Returns:       The error code for last error encountered on com1

(int).

NOTE: The setting for the lastError property is

cleared when errCount is reset or when the port is

successfully opened.

parity:        Use to set or get the desired parity scheme.

Arguments:     A number to indicate parity (int), where 0 = none

1 = odd and 2 = even.

Set Syntax:    r = com1:parity = 0

Returns:       If successful, r = 1; if not, r = error code

(int).

Get Syntax:    r = com1:parity

Returns:       The current parity scheme code for com1 (int).

Default:       0 (None) (int).

rxBufferSize:  Use to set or get the receive buffer size.

Arguments:     Any value between 2048 and 32768, inclusive (int).

Set Syntax:    r = com1:rxBufferSize = 2048

Returns:       If successful, r = 1; if not, r = error code

(int).

Get Syntax:    r = com1:rxBufferSize

Returns:       The current receive buffer size for com1 (int).

Default:       2048 (int).

rxTimeOutVal:  Use to set or get the receive timeout value, in

seconds.

Arguments:     Any value between 1 and 60, inclusive (int).

Set Syntax:    r = com1:rxTimeOutVal = 1

Returns:       If successful, r = 1; if not, r = error code

(int).

Get Syntax:    r = com1:rxTimeOutVal

Returns:       The current receive timeout value for com1 (int).

Default:       1 (int).

stopBits:      Use to set or get the number of stop bits.

Arguments:     1 or 2 (int).

Set Syntax:    r = com1:stopBits = 1

Returns:       If successful, r = 1; if not, r = error code

(int).

Get Syntax:    r = com1:stopBits

Returns:       The current number of stop bits for com1 (int).

Default:       1 (int).

txBufferSize:  Use to set or get the transmit buffer size.

Arguments:     Any value between 2048 and 32768, inclusive (int).

Set Syntax:    r = com1:txBufferSize = 2048

Returns:       If successful, r = 1; if not, r = error code

(int).

Get Syntax:    r = com1:txBufferSize

Returns:       The current transmit buffer size for com1 (int).

Default:       2048 (int).

txTimeDelay:   Use to set or get the transmit delay value, in

seconds.

Arguments:     Any value between 0 and 60, inclusive (int).

Set Syntax:    r = com1:txTimeDelay = 0

Returns:       If successful, r = 1; if not, r = error code

(int).

Get Syntax:    r = com1:txTimeDelay

Returns:       The current value for transmit delay between

characters for com1 (int).

Default:       0 (int).

verbose:       Use to enable or disable verbose operation.

Arguments:     1 to enable or 0 to disable (int).

Set Syntax:    com1:verbose = 1

Returns:       Always sets r = 1.

Get Syntax:    r = com1:verbose

Returns:       The current verbosity mode for com1 (int).

Default:       0 (Disabled) (int).

Methods:     Methods for the com1, com2, com3 and com4 objects are:

clearBreak:    Use to restart transmission and remove the port

from break state.

Arguments:     None.

Set Syntax:    r = com1:clearBreak

Returns:       If successful, r = 1; if not, r = error code

(int).

clearDTR:      Use to clear the DTR (data-terminal-ready) signal.

Arguments:     None.

Set Syntax:    r = com1:clearDTR

Returns:       If successful, r = 1; if not, r = error code

(int).

clearRTS:      Use to clear the RTS (request-to-send) signal.

Arguments:     None.

Set Syntax:    r = com1:clearRTS

Returns:       If successful, r = 1; if not, r = error code

(int).

close:         Use to close an open communications port.

Arguments:     1 (int).

Set Syntax:    r = com1:close = 1

Returns:       If successful, r = 1; if not, r = error code

(int).

errString:     Use to look up an error code and display its

explanation in the report area.

NOTE: The object used for this method is

irrelevant.

Arguments:     Error code (int).

Syntax:        string = com1:errString = -2

Returns:       The error string associated with the specified

error code.

flushRX:       Use to flush any characters in the receive queue.

Arguments:     None.

Get Syntax:    r = com1:flushRX

Returns:       If successful, r = 1; if not, r = error code

(int).

flushTX:       Use to flush any characters in the transmit queue.

Arguments:     None.

Get Syntax:    r = com1:flushTX

Returns:       If successful, r = 1; if not, r = error code

(int).

getRXcount:    Use to return the number of characters in the

receive queue.

Arguments:     None.

Get Syntax:    r = com1:getRXcount

Returns:       If successful, r = number of characters in the

queue; if not, r = error code (int).

getTXcount:    Use to return the number of characters in the

transmit queue.

Arguments:     None.

Get Syntax:    r = com1:getTXcount

Returns:       If successful, r = number of characters in the

queue; if not, r = error code (int).

helpCmds:      Use to display a brief description of object

methods in the report area.

NOTE: The object used for this method is

irrelevant.

Arguments:     None.

Syntax:        com1:helpCmds

Returns:       1 (int).

helpErrs:      Use to display brief a description of object

errors in the report area.

NOTE: The object used for this method is

irrelevant.

Arguments:     None.

Syntax:        com1:helpErrs

Returns:       1 (int).

Error Codes: NOTE: You can get an error code explanation using the

errString method, as described above.

Error codes for the com1, com2, com3 and com4 objects

are:

Code:  String:       Explanation:

-2     IE_OPEN       The requested port is already open.

-3     IE_NOPEN      The requested port is NOT open.

-4     IE_MEMORY     Memory allocation failure while creating

queues.

-5     IE_DEFAULT    Invalid port property was requested.

-10    IE_HARDWARE   A nonexistent port was requested.

-11    IE_BYTESIZE   An invalid number of data bits requested.

-12    IE_BAUDRATE   An invalid baud rate was requested.

-13    IE_CMDFAILED  The requested method (command) failed.

-14    IE_RXTIMEOUT  The port has timeout while receiving data.

-20    CE_RXOVER     A receive buffer overflow was detected.

-21    CE_OVERRUN    A receive character overrun was detected.

-22    CE_RXPARITY   A hardware parity error was detected.

-23    CE_FRAME      A hardware framing error was detected.

-24    CE_BREAK      A hardware break condition was detected.

-25    CE_CTSTO      CTS (clear-to-send) timeout during

transmission.

-26    CE_DSRTO      DSR (data-set-ready) timeout during

transmission.

-27    CE_RLSDTO     RLSD (receive-line-signal-detect) timeout

during transmission.

-28    CE_TXFULL     The transmit buffer is full.

-33    CE_MODE       An unsupported mode was requested.

helpProps:     Use to display brief description of object

properties in the report area.

NOTE: The object used for this method is

irrelevant.

Arguments:     None.

Syntax:        com1:helpProps

Returns:       1 (int).

msg:           Use to display brief port status message.

Arguments:     None.

Syntax:        com1:msg

Returns:       1 (int).

open:          Use to open the communications port.

Arguments:     1 (int).

Set Syntax:    r = com1:open = 1

Returns:       If successful, r = 1; if not, r = error code

(int).

Get Syntax:    r = com1:open

Returns:       The current port open status of com1 (int).

readBuffer:    Use to receive data into a buffer from the

selected port.

Arguments:     Number of characters to read (int).

Get Syntax:    buffer = com1:readBuffer = count

Returns:       If successful, buffer is valid; if not, buffer

= NULL.

readByte:      Use to receive data from the selected port.

Arguments:     None.

Get Syntax:    r = com1:readByte

Returns:       If successful, r = char; if not, r = error code

(int).

rpt:           Use to display complete port status in the report

window.

Arguments:     None.

Syntax:        com1:rpt

Returns:       1 (int).

setBreak:      Use to suspend transmission and place the port

in break state.

Arguments:     None.

Set Syntax:    r = com1:setBreak

Returns:       If successful, r = 1; if not, r = error code

(int).

setDTR:        Use to set the DTR (data-terminal-ready) signal.

Arguments:     None.

Set Syntax:    r = com1:setDTR

Returns:       If successful, r = 1; if not, r = error code

(int).

setRTS:        Use to set the RTS (request-to-send) signal.

Arguments:     None.

Set Syntax:    r = com1:setRTS

Returns:       If successful, r = 1; if not, r = error code

(int).

setXOFF:       Use to cause the transmission to act as if an XOFF

character was sent.

Arguments:     None.

Set Syntax:    r = com1:setXOFF

Returns:       If successful, r = 1; if not, r = error code

(int).

setXON:        Use to causes the transmission to act as if an XON

character was sent.

Arguments:     None.

Set Syntax:    r = com1:setXON

Returns:       If successful, r = 1; if not, r = error code

(int).

writeByte:     Use to transmit data out the selected port.

Arguments:     1 (int).

Set Syntax:    r = com1:writeByte = char

Returns:       If successful, r = 1; if not, r = error code

(int).

writeString:   Use to transmit a null-terminated string out the

selected port.

Arguments:     The string to transmit (string).

Set Syntax:    r = com1:writeString = string

Returns:       If successful, r = number of characters sent; if

not, r = error code (int).

Also See:    The example macro file named comsamp.mac and installed

into the x:\first\cfg\macros directory, where x: is the

drive on which you installed WinFIRST software.

###

___compare_____________________________________________________________

Description: Use to calculate a correlation coefficient by comparing

the current sample data with the current R1 reference,

or with a reference that is newly loaded (after you

specify it as ab argument).

Arguments:   (optional)  Name (including directory path, as needed)

of data file to compare with sample data (string).

Returns:     Correlation coefficient result (float).  The return

value is also placed in the z register.

Example:     load "pure1.abs"     /* Load file into memory. */

compare "pure2.abs"  /* Correlate sample with */

/* pure2.abs file. */

Also See:    compareDir

###

___compareDir__________________________________________________________

Description: Use to correlate sample data to all data files of the

same type in the specified directory.  CompareDir

searches the directory and displays the file names

and correlation coefficients of the 10 files most

similar to the sample data.

Arguments:   Name of directory containing data to compare (string).

Returns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:     load "pure1.abs"  /* Load file into memory. */

compare \winfirst\data\galaxy

/* Correlate all files in the */

/* galaxy directory. */

Also See:    compare

###

___concatenate_________________________________________________________

Description: Use to concatenate (join together) the sample data

file and the R1 (first) reference file.

NOTE: The resulting concatenated data file is placed

in a new data window.  (The original sample data and

reference data remain unchanged.)  This command is of

particular use in joining together multi-range spectra.

Arguments:   (optional)  The wavenumber values at which the two

files are to be joined.  If you do not specify a

point, WinFIRST automatically determines the point to

use by examining the starting and ending frequencies

of the two files.  A starting or ending value of the

sample or R1 reference is used as the concatenation

point.  Data after the concatenation point in the

sample file and data before the concatenation point

in the R1 reference file is discarded.

Returns:     1 for success, 0 for failure (int).

Example:     concatenate 1000 /* Join the sample and */

/* R1 reference data at */

/* 1000 wavenumbers. */

###

___copy________________________________________________________________

Description: Use to make and place a copy of a file in another

directory, or to make and place a copy of a file in the

same directory using a different file name.

Arguments:   One file name (string), and any of the symbols

and switches of the DOS COPY command.

Returns:     1 for success, 0 for failure (int).  (A random value is

placed in the z register.)

Example:     Copy "pure.abs" "\winfirst\backup\pure.abs"

/* Copy the file to the backup directory.*/

Also See:    cd, dir, erase, md, rename

###

___cos_________________________________________________________________

Description: Use to compute the cosine of the specified value.

Arguments:   Value (in radians) for which to find cosine (int or

float).

Returns:     Result (float).

Example:     y = cos 0  /* Place cosine of 0 in register y. */

x = cos z   /* Place cosine of value stored in */

/* register z into register x. */

Also See:    acos, asin, atan, sin, tan

###

___dataToArray_________________________________________________________

Description: Use to find the array position for a specified

wavenumber.

Arguments:   Wavenumber (float), array to use (variable).

Returns:     Wavenumber associated with position (float).

Example:     load co.abs  /* Load file into memory. */

x = dataToArray 2250.0 sample:data

/* Get array position for */

/* wavenumber 2250.0 and */

/* place value in register x. */

print x      /* Print the value to the screen. */

###

___date________________________________________________________________

Description: Date is an object that contains properties and methods

you can use to format and work with the system date.

Properties:  Properties for the date object are:

dateFormat:    Use to specify how the date is formatted when you

use the get method for the date object (date:get).

Arguments:     A text string that includes date format control

characters.  The default text string argument is

"%m%d%Y %H:%M:%S".  All available control

characters are as follows:

Character   Description

%a          Abbreviated weekday name

%A          Full weekday name

%b          Abbreviated month name

%B          Full month name

%c          Date and time representation

appropriate for the locale

%d          Day of the month as a decimal number

(01 - 31)

%H          Hour in 24-hour format (00 - 23)

%I          Hour in 12-hour format (01 - 12)

%j          Day of the year as a decimal number

(001 - 366)

%m          Month as a decimal number (01 - 12)

%M          Minute as a decimal number (00 - 59)

%p          Current locale's AM/PM indicator for a

12-hour clock

%S          Second as a decimal number (00 - 59)

%U          Week of the year as a decimal number,

with Sunday as the first day of the

week (00 - 51)

%w          Weekday as a decimal number (0 - 6,

Sunday is 0)

%W          Week of the year as a decimal number,

with Monday as the first day of the

week (00 - 51)

%x          Date represenation for current locale

%X          Time representation for current locale

%y          Year without the century as a decimal

number (00 - 99)

%Y          Year with the century as a decimal

number

%Z          Time zone name or abbreviation; no

characters if time zone is unknown

%%          Percent sign

Syntax:        date:dateFormat = "%m%d%Y %H:%M:%S"

Methods:     Methods for the date object are:

read:          Use to take a snapshot of the current system time.

Returns an integer that is the MS-DOS encoded

timestamp.

Arguments:     None.

Syntax:        date:read

get:           Use to return a string that is the time that was

last read with the date object's read method

(date:read), in the format specified by the date

object's dateFormat property (date:dateFormat).

Arguments:     None.

Syntax:        date:get

###

___dde_________________________________________________________________

Description: The dde object contains methods for Dynamic Data

Exchange (DDE).

NOTE: Refer also to the complete discussion of

WinFIRST software and dynamic data exchange in:

Appendix B  WinFIRST and Dynamic Data Exchange, in

the "WinFIRST User's Reference Guide."

DDE clients can establish three kinds of DDE links with

Mattson software acting as a server: cold links

(manual links), hot links (automatic links), and warm

links (notification links).  A DDE client makes a cold

link when it requests data one time only, or explicitly

requests data each time it is needed.  If the server's

data changes, the client connected via a cold link is

not aware of the change.  A hot link causes the server

to retransmit the data to the client when the data

changes.  This type of link is useful for applications

that want a continual stream of data.  The client does

not need to request the data directly.  Instead, the

Windows DDE manager keeps the hot link transaction

going.  A warm link is similar to a hot link, however

the server does not transmit changed data and instead

notifies the client when data has changed.  The client

must request that the changed data be transmitted if it

wants the changed data.

Methods for the dde object are:

advise:    Use to establish a constant data stream between

WinFIRST as it acts as a dde client, and all DDE

server applications.  The advise method instructs

WinFIRST to tell all servers with warm or hot links

to it to ask it for data updates.

NOTE: An alternate way to establish a constant data

stream into another application is to use WinFIRST

as a server and use the dde:poke method to send data

into the other application.  However, not all

applications provide DDE server support.

Arguments: None.

Syntax:    dde:advise

Returns:   1 for success, 0 for failure (int).

Example:   /* This example assumes another application has */

/* established a client-server relationship with *

/* "WINFIRST," "SYSTEM," and "A," */

/* where WINFIRST is the application */

/* name, SYSTEM is the topic name, and A is the */

/* item name (the contents of register A). */

A = 0

while (1) /* Loop until ctrl+break is pressed. */

{

bench:scan

A = peakAt 3150.0

dde:advise

doMessage /* Make the macro yield to Windows. */

}

execute:   Use to execute a command within the specified channel.

Arguments: Channel number (int), command (string).

Syntax:    dde:execute [channel number] [command]

Returns:   1 for success, 0 for failure (int).

Example:   r = dde:execute 1 "[FILE.CLOSE(FALSE)]"

initiate:  Use to initiate, or open an application for Dynamic

Data Exchange.

Arguments: Channel number, from 1 to 16 (int), application

(string), system topic (string).

Syntax:    dde:initiate [channel number] [application] [topic]

Returns:   1 for success, 0 for failure (int).

Example:   r = dde:initiate 1 "EXCEL" "SYSTEM"

poke:      Use to send data to the specified channel.

Arguments: Channel number, from 1 to 16 (int), item (string),

data (any value).  Allowed string values for item are:

DATA, REPORT, and any upper-case register (upper-case

A through upper-case Z).

Syntax:    dde:poke [channel number] [item] [value]

Returns:   1 for success, 0 for failure (int).

Example:   r = dde:poke 2 "DATA" 25.5

request:   Use to retrieve data from an application.

Arguments: Channel number, from 1 to 16 (int), item to retrieve

(string).  Allowed string values for item are: DATA,

REPORT, and any upper-case register (upper-case A

through upper-case Z).

Syntax:    dde:request [channel number] [item]

Returns:   Value requested for success, 0 for failure (int).

Example:   r = dde:request 2 "DATA"

terminate: Use to terminate an open channel.

Arguments: Channel number, from 1 to 16 (int).

Syntax:    dde:terminate [channel number]

Returns:   1 for success, 0 for failure (int).

Example:   dde:terminate 2

Also See:    In your "WinFIRST User's Reference Guide," "Appendix B

WinFIRST and Dynamic Data Exchange."

###

___decision loop (while loop)__________________________________________

Description: In general, use ApPro decision loops in a macro to ask

a yes-or-no question and to then continue according to

the resultant answer to the question.  When the answer

is "yes," the defined "while" block is executed, and

at the completion of the block, control returns to the

initial "while" statement.  If the user answers "yes"

to the question again, the "while" block is executed

again.  This continues until the user answers "no" to

the question.  See Chapter 2 in "WinFIRST Macros

and ApPro" for more information.

Example:     A = 0

while (ask "do you want to load a reference")

begin               /* Start of "while" block. */

loadref           /* Load a reference. */

end                 /* End of "while" block. */

Also See:    break, choice, choice and branch, for, if, while

Refer also to the expanded discussion of while loops

and other control statements in Chapter 2 of Using

"Using WinFIRST Macros and ApPro."

###

___decrement___________________________________________________________

Description: Use to decrease the value of a register by one.  This

command also works to drecrease an integer value that

is the last character in a file name.  (For example,

tmp3 becomes tmp2.)

Arguments:   Register to decrease (register).

Returns:     New value obtained in register (int).

Example:     float array usrarray[5]  /* Declare a floating point */

/* array that can hold 5 */

/* elements. */

x = 5          /* Set value of register x to 5. */

while (x > 0)  /* Loop until x is less than 0. */

begin          /* Mark start of loop. */

enter a "Please enter a wavenumber"

usrarray:data[x] = a  /* Assign the value stored */

/* in register a to the 5th */

/* position in the float */

/* arranamed 'usrarray'. */

decrement x  /* Decrease value of x by one. */

end            /* Mark end of loop. */

Also See:    increment

###

___deltaX______________________________________________________________

Description: Use to find the distance (in wavenumbers) between

each data point in the current sample.

Returns:     Result (float).  The return value is placed in the z

register.

Arguments:   None.

Example:     load "pure.abs"  /* Load file into memory. */

x = deltax       /* Get delta x and place in x. */

###

___deriv_______________________________________________________________

Description: Use to calculate the first derivative of the sample

data.  Aslo calculate higher order deriatives by using

the deriv command on the data repeatedly.

Arguments:   None.

Returns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:     load "pstyrene.abs"  /* Load file into memory. */

deriv  /* Calculate first derivative. */

show   /* Display results. */

###

___dir_________________________________________________________________

Description: Use to list the contents of a directory in a WinFIRST

report window.  This command works the same as the DOS

DIR command.

Arguments:   The same syntax and switches as the DOS DIR command.

Returns:     32 (int).  The return value is also placed in the z

register.

Example:     dir "\winfirst\data\galaxy"

/* List all files in the standard directory. */

Also See:    cd, copy, erase, md, rename

###

___display_____________________________________________________________

Description: Display is an object that contains properties you can

set and execute to specify how WinFIRST expand and

rescale operations work.

Properties:  Properties for the display object are:

annotLineColor: Use to specify the annotation line color.

Arguments:     24 bit color value.  Color values are specified

as triples of Blue, Green, and Red with 8 bits for

each color, red being the least significant byte

in the value and blue being the most significant.

Example color values (in hexadecimal) are:

Dark Blue   0x800000

Dark Red    0x000080

Dark Green  0x008000

Red         0x0000FF

Green       0x00FF00

Blue        0xFF0000

Light Gray  0xC0C0C0

Dark Gray   0x808080

Black       0x000000

White       0xFFFFFF

Cyan        0xFFFF00

Magenta     0xFF00FF

Yellow      0x00FFFF

Orange      0x0080FF

NOTE: Any value between 0 and 0xFFFFFF is

acceptable.  Windows will map the value to the

nearest color value.

Syntax:        display:annotLineColor=0x0080FF

bkgColor:      Use to specify the color of the data window.

Arguments:     24 bit color value.  Color values are specified

as triples of Blue, Green, and Red with 8 bits for

each color, red being the least significant byte

in the value and blue being the most significant.

Example color values (in hexadecimal) are:

Dark Blue   0x800000

Dark Red    0x000080

Dark Green  0x008000

Red         0x0000FF

Green       0x00FF00

Blue        0xFF0000

Light Gray  0xC0C0C0

Dark Gray   0x808080

Black       0x000000

White       0xFFFFFF

Cyan        0xFFFF00

Magenta     0xFF00FF

Yellow      0x00FFFF

Orange      0x0080FF

NOTE: Any value between 0 and 0xFFFFFF is

acceptable.  Windows will map the value to the

nearest color value.

Syntax:        display:bkgColor=0x0080FF

bottom:        Use to specify the bottom value of the y axis.

Arguments:     Floating point number (float).

Syntax:        display:top=1.0; display:bottom=0.0

box:           Use to specify the style of the inner frame.

Arguments:     The number that represents the style of the inner

frame or the string that represents the style of

the inner frame; where: 0 = "none", 1 = "2-sided",

2 = "4-sided", 4 = "bottom".

Syntax:        display:format = "none"

dataLineWidth: Use to specify the data line width to use for

plotting.

Arguments:      1 through 8 (int).

Syntax:        display:dataLineWidth = 1

expandXOnly:   Use to indicate whether or not you want only the

x axis to change during an expand procedure.

Arguments:     An index to identify how the expand procedure

works, where 0 = both the x and y axes expand and

1 = only the x axis is expanded during an expand

procedure.

Syntax:        display:expandXOnly = 1

format:        Use to specify the x axis style.

Arguments:     The number that represents the x axis style (int)

or the string that represents the x axis style

(string), where: 1 = "linear", 2 = "expanded",

4 = "microns".

Syntax:        display:format = "linear"

frame:         Use to specify the style of the outer frame.

Arguments:     The number that represents the style of the outer

frame (int) or the string that represents the

style of the outer frame (string), where:

0 = "none", 2 = "single", 4 = "double"

Syntax:        display:frame = "none"

frameColor:    Use to specify the color of frames.

Arguments:     24 bit color value.  Color values are specified

as triples of Blue, Green, and Red with 8 bits for

each color, red being the least significant byte

in the value and blue being the most significant.

Example color values (in hexadecimal) are:

Dark Blue   0x800000

Dark Red    0x000080

Dark Green  0x008000

Red         0x0000FF

Green       0x00FF00

Blue        0xFF0000

Light Gray  0xC0C0C0

Dark Gray   0x808080

Black       0x000000

White       0xFFFFFF

Cyan        0xFFFF00

Magenta     0xFF00FF

Yellow      0x00FFFF

Orange      0x0080FF

NOTE: Any value between 0 and 0xFFFFFF is

acceptable.  Windows will map the value to the

nearest color value.

Syntax:        display:frameColor=0x0080FF

gridColor:     Use to sp[ecify the grid color.

Arguments:     24 bit color value.  Color values are specified

as triples of Blue, Green, and Red with 8 bits for

each color, red being the least significant byte

in the value and blue being the most significant.

Example color values (in hexadecimal) are:

Dark Blue   0x800000

Dark Red    0x000080

Dark Green  0x008000

Red         0x0000FF

Green       0x00FF00

Blue        0xFF0000

Light Gray  0xC0C0C0

Dark Gray   0x808080

Black       0x000000

White       0xFFFFFF

Cyan        0xFFFF00

Magenta     0xFF00FF

Yellow      0x00FFFF

Orange      0x0080FF

NOTE: Any value between 0 and 0xFFFFFF is

acceptable.  Windows will map the value to the

nearest color value.

Syntax:        display:gridColor=0x0080FF

gridFormat:    Use to specify the grid format.

Arguments:     The number that represents the grid format (int)

or the string that represents the grid format

(string), where: 0 = "major" and

1 = "major+minor".

Syntax:        display:gridFormat = "major+minor"

NOTE: Major+minor is only applicable when both x

and y axes have minor tick marks.

gridStyle:     Use to specify the type of grid style.

Arguments:     The number that represents the grid style (int)

or the string that represents the grid style

(string), where: 0 = "none", 1 = "line",

2 = "cross" and 3 = "dot".

Syntax:        display:gridStyle = "line"

holeColor:     Use to specify the color of the data area.

Arguments:     24 bit color value.  Color values are specified

as triples of Blue, Green, and Red with 8 bits for

each color, red being the least significant byte

in the value and blue being the most significant.

Example color values (in hexadecimal) are:

Dark Blue   0x800000

Dark Red    0x000080

Dark Green  0x008000

Red         0x0000FF

Green       0x00FF00

Blue        0xFF0000

Light Gray  0xC0C0C0

Dark Gray   0x808080

Black       0x000000

White       0xFFFFFF

Cyan        0xFFFF00

Magenta     0xFF00FF

Yellow      0x00FFFF

Orange      0x0080FF

NOTE: Any value between 0 and 0xFFFFFF is

acceptable.  Windows will map the value to the

nearest color value.

Syntax:        display:holeColor=0x0080FF

left:          Use to specify the left value of the x axis.

Arguments:     Floating point number (float).

Syntax:        display:left=4000.0; display:right=400.0

refTitle:      Use to turn reference titles on or off.

Arguments:     0 or 1 (int) or "Off"/"On" (string).

Syntax:        display:refTitle = 0

rescaleToLimits:  Use to indicate how a click of the left mouse

in the data display area rescales the data

display.

Arguments:     An index to identify how the left mouse click

works, where 0 = the click autoscales the data

display in the usual method and 1 = the click

rescales using the current limits as specified on

the Display Limits dialog.  (Choose Set Limits on

the Data Area dialog to display the Display Limits

dialog.)

Syntax:        display:rescaleToLimits = 1

right:         Use to specify the right value of the x axis.

Arguments:     Floating point number (float)

Syntax:        display:left=4000.0; display:right=400.0

stackSize:     Use to specify the size of the reference stack.

Arguments:     Integer from 0 to 32 (int).

Syntax:        display:stackSize = 2

style:         Use to specify the display style.

Arguments:     The number that represents the display style to

use (int) or the string that represents the

display style (string), where: 1 = "stack",

2 = "overlay", 4 = "superimpose" and

8 = "sample only".

Syntax:        display:style = "stack"

tagger:        Use to turn display of the tagger on and off.

Arguments:     0 or 1 (int) or "Off"/"On" (string).

Syntax:        display:tagger = "On"

taggerString:  Use to specify the tagger control string.

Arguments:     The set of tagger information and formatting

codes that define the tagger you want to use

(string), where:

Tagger Information Codes

%t = sample file title

%f = sample file name and full directory path

%n = sample file base name (with no filename

extension)

%e = sample filename extension

%d = date and time of sample data collection

%o = operator name (as set in log book report)

%c = company name (as set in log book report)

%D = current time (updated as window is displayed

or redisplayed)

%r = resolution of sample data

%g = signal gain setting of sample data

%s = number of scans used to collect sample data

%i = iris setting used to collect sample data

%b = for Infinity spectrometers, the detector/

beamsplitter/source combination used to

collect the data

%1 = sample data file name and title, color coded

with sample data trace

%2 = first reference data file name and title,

color coded with reference data trace

%3 - %9 = file name and title, color coded with

reference data trace, for eight next

reference data files

Tagger Formatting Codes

^l = left justify

^c = center justify

^r = right justify

^t = tab (each tab is 12 character spaces)

^n = new line

Syntax:        display:taggerString = "%1^n%2^n%3"

tickLineWidth: Use to specify the frame and tick mark line

width to use for plotting.

Arguments:     1 through 8 (int).

Syntax:        display:tickLineWidth = 1

title:         Use to turn display of the data file title on and

off.

Arguments:     0 or 1 (int) or "Off"/"On" (string).

Syntax:        display:title = "On"

top:           Use to specify the top value of the y axis.

Arguments:     Floating point number (float).

Syntax:        display:top=1.0; display:bottom=0.0

Also See:    autoscale, autoscaleX, autoscale Y, expand, expandX,

expandY

###

___divide______________________________________________________________

Description: Use to divide the current sample data with the

data of another file.

Arguments:   The file name of the data by which you want to divide

the sample data in memory.

Returns:     1 for success, 0 for failure (int).  The return value

is placed in the z register.

Example:     load "pure1.abs"       /* Load file into memory. */

divide "pure2.abs"     /* Divide pure1.abs by data */

/* in the pure2.abs file. */

title "pure1 * pure2"  /* Assign title to results. */

show                   /* Display the results. */

Also See:    add, average, divide, multiply, subtract

###

___doMessage___________________________________________________________

Description: Use to allow Windows to process messages while WinFIRST

is running an ApPro macro.

NOTE: How the doMessage command works is modified in

WinFIRST version 3.0 and above. Previously, the

doMessage command waited for a Windows message to

process before returning control to the macro, and had

no return values.  Now, if a Windows message is

waiting, the message is processed, control is returned

to the macro, and the doMessage command returns a value

of 1.  But, if no message is found waiting, the

doMessage command immediately returns control to the

macro (without waiting for a message to process).  In

this case, the return value is 0.

The doMessage command is useful in macros that include

one or more loops that would otherwise, that is,

without the inclusion of the doMessage command, not

allow Windows to process messages.  For example,

including the doMessage command used in an Fkey loop

allows the user to still interact with Windows while

the macro is running.

Arguments:   None.

Returns:     1 if a Windows message was waiting and was processed,

0 if no Windows message was waiting (int).

Example:     closeallwindows

fkeys "1;2;3"

a=0

while (a!=3)

{

a = getfkey

if (a==1)

{

print "report"

print "line 1"

print "line 2"

report:maximize

}

else if (a==2)

{

b=1

while (b)

{

b=doMessage

}

}

}

NOTE: On some systems, and depending on what else is

running, the report window created when F1 is pressed

is not drawn correctly until F2 is pressed.

###

___dos_________________________________________________________________

Description: Use to execute a DOS internal command or batch file.

DOS internal commands are those that are loaded into

RAM when you first start the computer.  A batch file

contains a batch of DOS commands, which may be internal

or external commands.  For a complete description of

internal and external commands, as well as batch files,

please refer to your DOS manual.

Arguments:   The command or batch file name followed by its

required arguments.  These can be any valid argument

as long as they evaluate to a legal DOS command.

NOTE: If you use the ApPro dos command to execute the

DOS CD command, the resulting directory change is

temporary, and does not change your WinFIRST

environment.

Returns:     32 (int).  The return value is also placed in the

z register.

Example 1:   dos "type calib.dat"  /* Display the contents of */

/* the calib.dat file; TYPE */

/* is a DOS command. */

Example 2:   enter x "Enter the file to rename"

/* Ask for file to rename and store in x. */

enter y "Enter the new file name"

/* Ask for new file name and store in y. */

dos "rename" x y /* Rename the file whose name is */

/* stored in x using the name */

/* stored in y; RENAME is a DOS */

/* command. */

Also See:    execute, exit, launch, shell, system

###

___drawbox_____________________________________________________________

Description: Use to draw a box on the displayed data.

Arguments:   The x and y coordinates for a left corner of the box,

and for the right corner that is diagonal to the

left corner you specify.  The x coordinates are in

wavenumbers.  The y coordinates are in absorbance

units (all values are float).

Returns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:     load "pure.abs"            /* Load file into memory. */

drawbox 3500 0.3 2500 0.2  /* Draw a box at the */

/* specified location. */

Also See:    annotate, drawHruler, drawline, drawtext, drawVruler

###

___drawHruler__________________________________________________________

Description: Use to draw a horizontal ruler on the displayed data:

define two end points which are marked and labeled.

The absolute difference between the two end points is

calculated and labeled on the displayed ruler.

Arguments:   Two x-axis values, and one y-axis value (to locate the

horizontal ruler).

Returns:     The calculated absolute difference between the two

defined endpoints.  The return value is also placed in

the z register.

Example:     load "pure.abs"            /* Load file into memory. */

drawHruler 3500 2500 0.2   /* Draw a line at the */

/* specified location. */

Also See:    annotate, drawbox, drawline, drawtext, drawVruler

###

___drawline____________________________________________________________

Description: Use to draw a line on the displayed data.

Arguments:   The x and y coordinates for both the left and right

ends of the line.  The x coordinates are in

wavenumbers.  The y coordinates are in absorbance

units all values are float).

Returns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:     load "pure.abs"            /* Load file into memory. */

drawline 3500 0.3 2500 0.2 /* Draw a line at the */

/* specified location. */

Also See:    annotate, drawbox, drawHruler, drawtext, drawVruler

###

___drawtext____________________________________________________________

Description: Use to enter text (an annotation) on the displayed

sample data.

Arguments:   The x and y coordinates, with x in wavenumbers and

y as absorbance units (float), followed by the text

to display (string).

Returns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:     load "pstyrene.abs"  /* Load file into memory. */

show                 /* Display data. */

drawline 1601.02 0.796 2000 1.0    /* Draw line. */

drawtext 2300 1.1 "1601.02 0.796"  /* Draw label. */

Also See:    annotate, drawbox, , drawHruler, drawline, drawVruler

###

___drawVruler__________________________________________________________

Description: Use to draw a vertical ruler on the displayed data:

define two end points which are marked and labeled.

The absolute difference between the two end points is

calculated and labeled on the displayed ruler.

Arguments:   Two y-axis values, and one x-axis value (to locate the

vertical ruler).

Returns:     The calculated absolute difference between the two

defined endpoints.  The return value is also placed in

the z register.

Example:     load "pure.abs"            /* Load file into memory. */

drawVruler .2 .4 3500      /* Draw a line at the */

/* specified location. */

Also See:    annotate, drawbox, drawHruler, drawline, drawtext

###

___dx2matt_____________________________________________________________

Description: Use to a translate a copy of a file in the JCAMP-DX

format to a file in the Mattson FIRST data file

format.  The translated file becomes the current

sample.  Four data types (stored in the JCAMP format)

can be translated: absorbance, transmittance, diffuse

reflectance, and single beam.  The original JCAMP file

that is translated is not removed or otherwise altered

by the dx2matt command.  You can use the translated

data file in the same way you would any data file

created by WinFIRST.

Arguments:   The directory path and name of JCAMP-DX file (string).

Returns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:     dx2matt "\JCAMP\data\pstyrene"

/* Translate file and make it the */

/* sample data file. */

###

___enter_______________________________________________________________

Description: Use to 1) display a dialog that prompts the user to

enter a value (integer, float, or text string) and to

2) place the value the user enters into a register you

specify as an argument for the enter command.  You also

specify the prompting text that is displayed.  And, the

current value of the register you specify is always

displayed as the "default" response to the prompt.

Arguments:   A register and the text of the question or directions

to display (string).

Example:     enter n "Enter file name to load into memory"

/* Ask user for file name and place in n. */

load n  /* Load file specified by n into memory. */

###

___enterString_________________________________________________________

Description: Use to display a dialog that prompts the user to

enter a text string as a response.  The return is

always a string and can be placed in a register or

a string variable.  This command differs from the

enter command which can use only a register.

Arguments:   (optional)  The first argument is the text of the string

to display as the prompt on the dialog (string).  If

none is provided, then "->" is used as the default.

The second argument is the text of the string to

display as the default response in the edit control

(string).  If a second argument is not provided, no

default string appears in the edit control.  The

maximum number of characters you can enter is 255.

Returns:     The response string.

Example:     enterString "prompt" "default"

###

___eof_________________________________________________________________

Description: Use to tell if you are at the end of a file that is

open for input and output.

Returns:     1 if at end of file, and 0 if not (int).  Also, the

return value is stored in the z register.

Arguments:   Number of the file.

Example:     while (!(eof a))   /* While not at end of file a... */

begin

fread 1 a        /* Read value to register a. */

print a

end

Also See:    append, exist, fclose, filepos, finput, fopen, fprint,

fread, freadByte, fseek, fwrite, fwriteByte, rewind

###

___erase_______________________________________________________________

Description: Use to erase (delete) a file.

NOTE: Use erase with care.  Once you erase a file, you

cannot retrieve it.

NOTE: Wildcard characters, such as * or ?, are not

allowed.

Arguments:   Name of file you want to erase (string).

Returns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:     erase red.bkg  /* Erase the file. */

Also See:    cd, copy, dir, md, rename

###

___errorMsg____________________________________________________________

Description: Use to display an error message or a yes or no

question at the ApPro command line.

Arguments:   Any valid ApPro values: register, variable, or literal.

Returns:     1 if y is typed, 0 if not (int).  The return value is

also placed in the z register.

Example:     fopen 1 sample.txt  /* Open sample.txt as file 1. */

fread 1 a           /* Read value to register a. */

eof 1               /* Find out if at end of file 1. */

if (z == 1)         /* If at file end... */

errorMsg "End of file reached...Continue?"

/* Display error message and */

/* ask whether to continue. */

if (z == 0)         /* If y is not typed... */

exit              /* Exit the macro.  Otherwise, */

/* continue the macro. */

###

___escapeHit___________________________________________________________

Description: Use to determine if the Esc key was pressed during

execution of the macro.  This command is useful in

cases where a miscalculation in the code causes an

infinite loop.  If the loop contains code to check for

the Esc key, you can terminate the loop.

NOTE: The CTRL + BREAK key combination terminates the

entire macro while the escapehit command gives you

another (and sometimes preferable) option.

Arguments:   None.

Returns:     Non-zero if Esc is pressed, otherwise 0.

Example:     flushkb        /* Clear the keyboard buffer. */

while (x > 0)  /* Loop while x is greater than 0. */

begin          /* Mark start of loop. */

y = y + x      /* Add value of x to y. */

decrement x    /* Decrease the value of x by 1. */

if (escapehit) /* If Esc key is pressed... */

break        /* Stop the execution of this loop. */

end            /* Mark end of loop. */

###

___execString__________________________________________________________

Description: Use to execute a text string as ApPro commands.

Arguments:   Text string to execute. (string)

Returns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:     /* This macro gives information for an array: */

string arrname  /* Declare a string to hold */

/* the array name. */

string field    /* Declare a string to hold */

/* the array field. */

int done        /* Declare an integer to control */

/* the loop. */

done = 0        /* Set done equal to 0. */

enter a "Please enter the name of the array"

/* Ask user and store response in a */

arrname = a     /* Set arrname equal to the text */

/* stored in a. */

/* This macro prints a list of the valid fields to */

/* access: */

print "valid fields are:"

print "firstx, lastx, start, rows, cols, pages, and "

print "data[ array pos ] ([][])"

while (!done)  /* While done is non-zero. */

{              /* Mark start of loop. */

enter a "Enter array field or Q to quit"

if (stringcompare a "Q")  /* If Q is entered... */

done = 1                /* Set done to 1. */

else       /* If the string compare fails... */

{          /* Mark start of else statements. */

field = a  /* Assign text in a to the field */

/* variable. */

execstring (buildstring "print " arrname ":" field)

/* If arrname is "sample" and field is */

/* "firstx", then the string executed is */

/* 'print sample:firstx'. */

}        /* Mark end of else statements. */

}          /* Mark end of loop. */

Also See:    buildString, getStringElement, isNumber,

setStringElement, string2Float, string2Int,

stringCompare, stringIndex, stringLeft, stringLength,

stringLower, stringMid, stringRight, stringUpper

###

___execute_____________________________________________________________

Description: Use to run and to pass control to any DOS or Windows

executable program.  As control is passed to the

DOS or Windows program, the WinFIRST software

is terminated (quits running).

Arguments:   Any valid command from the DOS command line (string).

Returns:     None, as WinFIRST is terminated.

Example:     execute "edlin test.fil"

/* Execute edlin for test.fil. */

Also See:    dos, exit, launch, shell, system

###

___exist_______________________________________________________________

Description: Use to determine whether a file exists in the current

directory, or a specified directory.

Arguments:   file name or full path and file name (string).

Returns:     1 if file exists, 0 if not (int).  Also, the return

value is stored in the z register.

Example:     exist "\winfirst\data\galaxy\pure.abs"

/* Check to see if file pure.abs exists. */

Also See:    append, eof, fclose, filepos, finput, fopen, fprint,

fread, freadByte, fseek, fwrite, fwriteByte, rewind

###

___exit________________________________________________________________

Description: Use for two purposes: 1) to exit a DOS shell (see the

shell command) and return to WinFIRST and ApPro, and

2) to exit an ApPro macro or ApPro command line.

Arguments:   None.

Returns:     None.

Example:     shell  /* Create a temporary copy of DOS.  The DOS */

/* prompt appears. */

exit   /* Close the DOS copy (shell) and return to */

/* WinFIRST ApPro. */

Also See:    dos, execute, launch, shell, system

###

___exp_________________________________________________________________

Description: Use to calculate e to the x for a given value, or

for the sample data if no value is given.

Arguments:   (optional)  Value to use to compute e to the x (float).

Returns:     Result if a value is specified (float), 1 if a value is

not specified (int), or 0 for failure (int).

Example 1:   load "pstyrene.abs"  /* Load file into memory. */

exp                  /* Calculate e to the x for */

/* each data point in the */

/* sample data. */

show                 /* Show results. */

Example 2:   print (exp 2.5)      /* Calculate e to the 2.5. */

Also See:    ln, log

###

___expand______________________________________________________________

Description: Use to expand the display of a region of the

sample data.

NOTE: Use the autoscale command to restore the data

display to the unexpanded display.

Arguments:   Two wavenumbers and two y axis values, as follows:

leftmost wavenumber, rightmost wavenumber, top y

value, bottom y value (float).

Returns:     1 (int).  The return value is also placed in the z

register.

Example:     load "pstyrene.abs"       /* Load file into memory. */

expand 3500 2500 2.5 0.0  /* Expand region. */

show                      /* Show expanded display. */

autoscale                 /* Show as unexpanded. */

Also See:    autoscale, autoscaleX, autoscaleY, display:expandXOnly,

display:rescaleToLimits, expandX, expandY

###

___expandX_____________________________________________________________

Description: Use to show an enlarged or expanded region of the

SAMPLE data along the x axis.

NOTE: Use the autoscale command to restore the data

display to the unexpanded display.

Arguments:   Two wavenumbers, as follows: leftmost wavenumber,

rightmost wavenumber (float).

Returns:     1 (int).  The return value is also placed in the z

register.

Example:     load "pstyrene.abs"  /* Load file into memory. */

expandx 3500 2500    /* Expand display in x region. */

show                 /* Show expanded display. */

autoscale            /* Show unexpanded display. */

Also See:    autoscale, autoscaleX, autoscaleY, display:expandXOnly,

display:rescaleToLimits, expand, expandY

###

___expandY_____________________________________________________________

Description: Use to show an expanded region of the sample data

along the y axis.

NOTE: Use the autoscale command to restore the data

display to the unexpanded display.

Arguments:   Two y axis values, as follows: top y value, bottom y

value (float).

Returns:     1 (int).  The return value is also placed in the z

register.

Example:     load "pstyrene.abs"  /* Load file into memory. */

expandy 2.0 0.0      /* Expand display in y region. */

show                 /* Show expanded display. */

autoscale            /* Show unexpanded display. */

Also See:    autoscale, autoscaleX, autoscaleY, display:expandXOnly,

display:rescaleToLimits, expand, expandX

###

___extension___________________________________________________________

Description: Use to return the filename extension, without the "."

separator, of a file name you specify.

Arguments:   The file name (filename plus filename extension).

Returns:     The filename extension of the specified file name,

without the "." separator.

Example:     x = extention "pstyrene.abs"

###

___fclose______________________________________________________________

Description: Use to close a data file that is open for read and

write.

Arguments:   File number (int).

Returns:     1 for success, 0 for failure (int).  (A random value

is placed in the z register.)

Example:     fopen 1 "file1.dat"  /* Open file1.dat for read and */

/* write. */

fread 1 a            /* Read first value in file 1*/

/* and store in register a. */

fclose 1             /* Close file number 1. */

Also See:    append, eof, exist, filepos, finput, fopen, fprint,

fread, freadByte, fseek, fwrite, fwriteByte, rewind

###

___fft_________________________________________________________________

Description: fft is an object that contains properties and methods

you can set and execute to perform Fast Fourier

Transform processing of interferogram data.

NOTE: Enclose all strings entered as arguments within

double quotation marks.

For example:

load test.igm    /* Load interferogram file. */

fft:setTodata    /* As a needed first step, set */

/* all current fft process */

/* properties to match those used */

/* to process active sample data. */

/* This step assures properties, */

/* like ndata, symmetry, etc. */

/* have valid settings.  */

fft:ftApod = "triangle"  /* Set properties as */

/* needed. */

fft:process      /* Process interferogram data. */

Properties:  Properties for the fft object are:

fend:          Use to specify a wavenumber for the end of the

frequency range of the processed data file.

Arguments:     Enter a value (int or string) appropriate for the

spectral range that the input data represents:

For NIR, the default starting value is 2000

wavenumbers.  Enter a value from 1 to 15,798.

For MIR, the default starting value is 400

wavenumbers.  Enter a value from 1 to 7,899.

For FIR, the default starting value is 50

wavenumbers.  Enter a value from 1 to 1,974.

For VIS, the default starting value is 15,798

wavenumbers.  Enter a value from 15,798 to 31,596.

For UV, the default starting value is 31,596.

Enter a value from 31,596 to 47,393.

Syntax:        fft:fend = 4000

fstart:        Use to specify a wavenumber for the start of the

frequency range of the processed data file.

Arguments:     Enter a value (int or string) appropriate for the

spectral range that the input data represents:

For NIR, the default starting value is 2000

wavenumbers.  Enter a value from 1 to 15,798.

For MIR, the default starting value is 400

wavenumbers.  Enter a value from 1 to 7,899.

For FIR, the default starting value is 50

wavenumbers.  Enter a value from 1 to 1,974.

For VIS, the default starting value is 15,798

wavenumbers.  Enter a value from 15,798 to 31,596.

For UV, the default starting value is 31,596.

Enter a value from 31,596 to 47,393.

Syntax:        fft:fstart = 400

ftApod:        Use to specify the apodization method to use.

Arguments:     The name for the type of apodization method you

want to use: boxcar, triangle, Happ Genzel,

Bessel, Welch, Beer Norton Weak, Gaussian, Beer

Norton Medium, Beer Norton Strong, Hanning, Cosine

Squared, Cosine 4th, or 1-Cosine 8th (string).

Syntax:        fft:ftapod = "triangle"

ftOutput:      Use to specify the output form of the processed

data.

Arguments:     The name of the output form you want the FFT

processing to create: single beam, transmittance,

absorbance, or diffuse reflectance (string).

Syntax:        fft:ftOutput = "single beam"

ftPhaseApod:   Use to specify the phase apodization method.

Arguments:     The name for the type of phase apodization you

want to use: boxcar, triangle, Happ Genzel,

Bessel, Welch, Beer Norton Weak, Gaussian, Beer

Norton Medium, Beer Norton Strong, Hanning, Cosine

Squared, Cosine 4th, or 1-Cosine 8th (string).

Syntax:        fft:ftPhaseApod = "Happ Genzel"

ftPhaseMethod: Use to specify the phase correction method.

Arguments:     The name for the type of phase correction method

you want to use: Mertz, Stored, Power, or None

(string).

Syntax:        fft:ftPhaseMethod = "power"

ftPhaseType:   Use to specify the phase type used to perform

the FFT.

Arguments:     The name for the type of phase type you want to

use: imaginary or real (string).

Syntax:        fft:ftPhaseType = "imaginary"

ftSymmetry:    Use to specify FFT symmetry as either double or

single.

Arguments:     The name (string) for the type of symmetry you

want to use: single or double.

Syntax:        fft:ftsymmetry = "double"

ndata:         Use to set the number of data points to use in the

FFT calculation (the number of interferogram data

points to transform).  This value (along with

lasersamp and FFT Symmetry) determines the

resulting spectrum's optical resolution.  Notice

that ndata can be different than nfft.

NOTE: Set the ndata property only if you are

knowledgeable in the area of FFT processing.

Arguments:     Any value (int or string).

Syntax:        fft:ndata = 1024

nfft:          Use to specify the number of Fast Fourier

Transform points, that is, the total number of

data points to transform.  The number of FFT

points determines digital resolution, defined as

the distance between output data points.

Arguments:     A value (int or string) to specify the number of

FFT points.  This number must be a power of 2, no

smaller than 1024 and no greater than 512K, or

524288.

Syntax:        fft:nfft = 1024

nPhaseData:    Use to specify the number of phase data points.

Arguments:     A value (int or string) to specify the number of

phase data points.

Syntax:        fft:nPhaseData = 300

PhaseNfft:     Use to specify the number of phase FFT points.

(For example, if overshoot, or ringing appears at

the edge of a strong, wide band, larger values for

PhaseNfft, and also nPhaseData may be needed.)

Arguments:     A value (int or string) to specify the number of

phase FFT points.  This number must be a power of

2, no less than 64 and no more than 512K, or

524288.

Syntax:        fft:PhaseNfft = 1024

refSource:     Use to specify the source for the background data

file to use in processing the interferogram data.

Arguments:     The name for the background data source you want

to use in FFT processing: Current Background, Disk

File (File Open dialog prompts you to select

file), or Current Window (string).

Syntax:        fft:refSource = "current background"

Methods:     Methods for fft are:

makeBackground: Use to make the processed data the current

background.

Arguments:     None.

Syntax:        fft:makeBackground

process:       Use to perform the FFT.

Arguments:     None.

Syntax:        fft:process

readMethod:    Use to load a math method (configuration).  The

fft property settings specified by the math method

are put in place as the current settings.

Arguments:     The name of the method file to load (string).

Syntax:        fft:readmethod = "forfft"

saveMethod:    Use to save a math method (configuration).  The

current fft property settings are saved as part of

the method.

Arguments:     The name of the method file to save (string).

Syntax:        fft:savemethod

setToData:     Use to set current fft properties to match those

used to process the current sample data file (as

specified in data file header).  Modify individual

property settings as needed.

Arguments:     None.

Syntax:        fft:settodata

###

___filePos_____________________________________________________________

Description: Use to determine the current position of a text or

binary file that is open for read and write.  The file

position is given in number of bytes from the beginning

of the file.

Arguments:   File number (int).

Returns:     The position, in bytes (int).  The return value is also

placed in the z register.

Example:     fopen 1 "file1.dat"  /* Open file 1 for read and */

/* write. */

print (filepos 1)    /* Print file position for */

/* file 1. */

Also See:    append, eof, exist, fclose, finput, fopen, fprint,

fread, freadByte, fseek, fwrite, fwriteByte, rewind

###

___fileTime____________________________________________________________

Description: Use to determine the time a file was created.  Use

the fileTime command in combination with the clock

command to determine when a file was created.  To set

a timer to keep track of seconds, use the timerCheck

and timerStart commands.

Arguments:   The name of the file whose creation time you want to

determine. (string)

Returns:     The time, in seconds (int).  The return value is also

placed in the z register.

Example:     /* This program determines the age of a file.  Then, */

/* if the file is more than a week old, the program */

/* calls a user-defined function to acquire another */

/* sample: */

int age          /* Declare integer for file age. */

int current      /* Declare integer for current */

/* time. */

current = clock  /* Get the current time and store */

/* the value in current. */

age = filetime "tmp.abs"  /* Get the creation time */

/* of the tmp.abs file. */

if (( current - age )  > 648000)  /* If too old... */

getnewsamp      /* Call this user-defined function. */

/* If file is less than 1 week old */

/* (604800 seconds = 1 week), do */

/* nothing. */

###

___fillArea____________________________________________________________

Description: Use to fill a specified wavenumber range to a

calculated baseline between two wavenumbers.

Arguments:   Two wavenumbers (float).

Returns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:     load "pstyrene.abs"    /* Load file into memory. */

fillarea 3000 2500     /* Fill wavenumber region to */

/* calculated baseline. */

show                   /* Show the results. */

###

___findFirstFile_______________________________________________________

Description: Use to find the first occurrence of a set of files with

similar names within the current directory.

Arguments:   Any valid DOS filename with optional wildcards, ? and

* (string).

Returns:     The name of the first file found, or an empty (blank)

string if nothing is found (string).

Example:     /* This prints the names of all files in the current */

/* directory that have .mac as a filename extension: */

s = findfirstfile "*.mac"  /* Put the first filename */

/* found into register s. */

a = stringlength s         /* Put length of string */

/* in s into register a. */

while ( a > 0 )     /* Loop until a is 0 length. */

{                   /* Mark start of loop. */

print s           /* Print file name to screen. */

s = findnextfile  /* Get next occurrence of *.mac. */

a = stringlength s  /* Put string length into a. */

}                   /* Mark end of loop. */

###

___findNextFile________________________________________________________

Description: Use to find the next occurrence of the set of files

specified by the findFirstFile command.

Arguments:   None.

Returns:     The name of the file found, or an empty (blank) string

if nothing is found (string).

Example:     /* This prints the names of all files within the */

/* current directory that have .mac as a filename */

/* extension: */

s = findfirstfile "*.mac"  /* Put the first filename */

/* found into register s. */

a = stringlength s         /* Put length of string */

/* in s into register a. */

while ( a > 0 )     /* Loop until a is 0 length. */

{                   /* Mark start of loop. */

print s           /* Print file name to screen. */

s = findnextfile  /* Get next occurrence of *.mac. */

a = stringlength s  /* Put string length into a. */

}                   /* Mark end of loop. */

###

___finput______________________________________________________________

Description: Use to read one line of text from a text file.

Arguments:   The file number of an open file (int), and a register

in which to place the read line (register).

Example:     fopen 1 "testfile.dat"  /* Open testfile.dat file. */

finput 1 a              /* Read first line in file. */

finput 1 b              /* Read next line in file. */

fclose 1                /* Close the file. */

print "line 1  :  " a   /* Print the contents of */

/* register a to the screen. */

print " line2  :  " b   /* Print the contents of */

/* register b to the screen. */

Also See:    append, eof, exist, fclose, filepos, fopen, fprint,

fread, freadByte, fseek, fwrite, fwriteByte, rewind

###

___fkeys_______________________________________________________________

Description: Use to define a row of from one to eight function keys

to display along the bottom of the WinFIRST

application window.  Use the fkeys command in

combination with the getFkey command to 1) pause the

macro with the function keys displayed, 2) to display a

prompt on the status line below the function keys, and

3) to retreive user input (to recognize which function

key the user selected).

Arguments:   A list of up to eight labels you want to assign to

function keys (string).  Labels must be separated

from each other within the list by a semi-colon (;).

Returns:     2 for a valid label list, 1 for an empty label list

(int).  The return value is also placed in the z

register.

Example:     int choice  /* Declare integer variable to hold key */

/* selection. */

int done    /* Declare integer variable for loop */

/* control. */

done = 0    /* Set done equal to 0. */

fkeys "Sample; Bkg; Plot; Done"

/* Define function keys. */

while (done == 0)  /* Loop until done is non-zero. */

{                  /* Mark start of loop. */

choice = getFkey "Press a function key..."

/* Wait for user to press a function */

/* key.  After a key is pressed, the */

/* key number is stored in choice. */

if (choice == 1)   /* If first Fkey pressed... */

collectsample    /* Call this user function. */

else if (choice == 2) /* If second Fkey pressed... */

collectbkg       /* Call this user function. */

else if (choice == 3) /* If third Fkey pressed... */

plotcurrentdata     /* Call this user function. */

else if (choice == 4) /* If fourth Fkey pressed... */

done = 1         /* Set done to 1 to terminate */

/* the loop. */

}                    /* Mark end of while loop. */

Also See:    fkeysOff, getFkey, menu

###

___fkeysOff____________________________________________________________

Description: Use to turn off the display of the set of function keys

that the macro previously defined.

Arguments:   None.

Returns:     1 for success, 0 for failure (int).

Example:     /* This macro provides one way to give */

/* instructions to the user within a larger */

/* macro and pause while the user reads the */

/* instructions. (Another way is to use the */

/* built-in macro command named pause.  The */

/* pause command puts up a message box that may */

/* be inconveniently on top of a report window. */

/* While the pause message box is movable, it */

/* may sometimes be more convenient to use the */

/* pause-and-message-display technique */

/* illustrated here.  This macro uses the fkeys */

/* and fkeysOff commands to place the message */

/* prompt on the status line. */

report:new       /* Open a new report window. */

print "Place instructions for the user in the "

print "new report window using the print "

print "command.  Or, use the showFile command "

print "to display a text file of instructions."

print " "

print "Press F1 to continue."

report:maximize  /* Maximize the report window. */

fkeys "Press F1 to continue."

while ( 1)

{

A = getFkey "Press F1 to continue."

if (A == 1)

{

break

}

}

fkeysOff

report:close     /* Close the report window. */

Also See:    fkeys, getFkey, menu, pause

###

___fLength_____________________________________________________________

Description: Use to report the integer length of the specified file.

Arguments:   The name of the file for which you want to report

integer length.

Returns:     The integer length for success, 0 for failure (int).

The return value is also placed in the z register.

Example:     a = flength "filename" /* Place the file's integer */

/* length in the a register. */

print a                /* Display the integer value */

/* in the WinFIRST report */

/* window. */

Also See:    fSum

###

___flushKb_____________________________________________________________

Description: Use to clear the keyboard buffer when running a macro.

For example, use this command when using the escapeHit

command.  (After using flushKb, you are sure the

keyboard buffer is empty.  Then when the Esc key is

pressed, it is the first key in the buffer.)

Arguments:   None.

Returns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:     flushkb         /* Clear the keyboard buffer. */

while (x > 0)   /* Loop while x is greater than 0. */

begin           /* Mark start of loop. */

y = y + x       /* Add value of x to y. */

decrement x     /* Decrease value of x by 1. */

if (escapehit)  /* If Esc key pressed... */

break;        /* Stop executing this loop. */

end             /* Mark end of loop. */

###

___fopen_______________________________________________________________

Description: Use to open a file for reading and writing.  Up to 16

files can be open at one time.

Arguments:   File number, from 1 to 16 (int) and file name (string).

Returns:     5 (int).  The return value is also placed in the z

register.

Example:     fopen 1 "file1.dat"  /* Open file1.dat for read */

/* and write. */

fread 1 a            /* Read first value in file 1 */

/* into register a. */

fclose 1             /* Close file number 1. */

Also See:    append, eof, exist, fclose, filepos, finput, fprint,

fread, freadByte, fseek, fwrite, fwriteByte, rewind

###

___for_________________________________________________________________

Description: Use the for command to define a for loop, a special

ApPro instruction that repeats a statement list a set

number of times.  How many times the statement list

repeats is determined by changing the value of a

control variable.  The control variable can be any

user-defined variable of the int or float type or any

register.

NOTE: Use of a begin-end statement is always required

in a for loop.

NOTE: Comments are not allowed alone on a line anywhere

within a for loop, with one exception: a comment is

allowed alone on a line as long as the comment is

contained within a begin-end statement.  As part of the

for loop instruction, you specify both the starting and

the ending value of the control variable.  You also

specify a step value that is the amount of increase (or

decrease) to apply to the control variable with each

execution of the statement list.

Arguments:   None.

Syntax:      for <variable or register> = exp1 to exp2 step exp3

begin

statement list

end

where:

exp1  is the first value to assign to the control

variable.

exp2  is the last value to assign to the control

variable that still allows execution of the loop.

exp3  is the step value, that is, the amount by which

the control variable increases (or decreases).

Example:     for x = 0 to 9 step 1

begin

print "x = " x

print "x squared = " (x * x)

end

Also See:    break, choice, choice and branch, decision loop, if,

while

Refer also to the expanded discussion of for loops and

other control statements in Chapter 2 of Using WinFIRST

Macros and ApPro.

###

___fprint______________________________________________________________

Description: Use to print text to an open (text) file.

Arguments:   File number, from 1 to 16 (int) and text to print to

file (up to 19 literal values, variables or registers).

Returns:     1 for success, 0 for failure (int).

Example:     fopen 1 "file1.dat"  /* Open file1.dat for read */

/* and write. */

a = "file1.dat"      /* Set a equal to "file.dat". */

fprint 1  "Filename :  " a  /* Print to the file. */

fclose 1             /* Close file number 1. */

Also See:    append, eof, exist, fclose, filepos, finput, fopen,

fread, freadByte, fseek, fwrite, fwriteByte, rewind

###

___frame_______________________________________________________________

Description: Frame is an object with methods you can use to size

and can use to move and size the WinFIRST application

(frame) window.

NOTE: It is possible to position the WinFIRST

application window off the screen, which has the effect

of hiding the application.

NOTE: It is usually necessary to give a DoMessage

command after a frame command to allow Windows an

opportunity to service the request.

Methods:     Methods for the frame object are:

hide:          Use to hide the WinFIRST application window, for

example, while running a macro.

NOTE: After using frame:hide, to hide WinFIRST, be

sure to use a subsequent frame:show, to show

WinFIRST.  If you do not, WinFIRST will continue

to run, but you will have no way to get to it.

(Windows does not include "invisible" application

on its Task List.  Similarly, Windows does not

present "invisible" applications when you use the

Alt-Tab key combination to choose amoung open

applications.)

Arguments:     None.

Syntax:        frame:hide

maximize:      Use to maximize the WinFIRST application window in

your display screen.

Arguments:     None.

Syntax:        frame:maximize

minimize:      Use to minimize the WinFIRST application window to

an icon.

NOTE: If when you use the frame:minimize command

to minimize WinFIRST either the WinFIRST macro

toolbox or the WinFIRST control panel, or both,

are displayed, these WinFIRST tools remain

displayed.  This means you can run macros and

collect data without even seeing the WinFIRST

application window.

Arguments:     None.

Syntax:        frame:minimize

move:          Use to move the WinFIRST application window.

Arguments:     Four values to show, in order, 1) the top left

corner x-axis position, 2) the top left corner

y-axis position, 3) the frame width, and 4) the

frame height.  All values are in pixels of monitor

resolution.  For example, standard VGA monitor

resolution is 640 by 480 pixels.

Syntax:        frame:move 0 0 100 200

show:          Use to show the WinFIRST application window when

it is hidden, for example, while running a macro.

Arguments:     None.

Syntax:        frame:show

###

___fread_______________________________________________________________

Description: Use to read 8-byte double precision floating point

binary data from an open user file.  One value is read

at a time.

Arguments:   The file number (from 1 to 16) to read from (int) and

the register into which to place the read value.

Returns:     1 for success, 0 for failure (int).  And, the read

value is placed in the specified register.

Example:     fopen 1 "file1.dat"  /* Open file1.dat for read */

/* and write. */

fread 1 a            /* Read first value in file 1 */

/* into register a. */

fclose 1             /* Close file number 1. */

Also See:    append, eof, exist, fclose, filepos, finput, fopen,

fprint, freadByte, fseek, fwrite, fwriteByte, rewind

###

___freadByte___________________________________________________________

Description: Use to read raw ASCII data from a file.

Arguments:   File number, from 1 to 16 (int).

Returns:     Returns the value read.  Also, places in the z register

1 for success, or 0 for failure (int).

Example:     if (fopen 1 "c:\bytes.bin")

begin

a = 20

b = 40

c = 80

print "write " a " " b " " c

fwritebyte 1 a b c

print "write 20 30 40"

fwritebyte 1 20 30 40

rewind 1

a = freadbyte 1

b = freadbyte 1

c = freadbyte 1

print "read " a " " b " " c

a = freadbyte 1

b = freadbyte 1

c = freadbyte 1

print "read " a " " b " " c

fseek 1 3

a = freadbyte 1

print "read byte 3: " a<R><T>fclose 1

end

NOTE: The byte read after the fseek command is "20,"

because the byte counting starts at 0, not at 1.  That

is, the first byte in the file is at offset 0, the

second byte is at offset 1, etc.  Thus, "fseek 1 3"

moves the file pointer to offset 3 in the file, which

contains the fourth byte.

Also See:    append, eof, exist, fclose, filepos, finput, fopen,

fprint, fread, fseek, fwrite, fwriteByte, rewind

###

___fseek_______________________________________________________________

Description: Use to move to a specific position in a text or binary

file that is open for read and write.

Arguments:   File number (int), number of bytes to move into the

file (int).

Returns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:     fopen 1 "file.dat"  /* Open file.dat for read and */

/* write. */

fseek 1  32         /* Seek 32 bytes into the file. */

fread 1 a           /* Read the next value in file. */

fclose 1            /* Close the file. */

Also See:    append, eof, exist, fclose, filepos, finput, fopen,

fprint, fread, freadByte, fwrite, fwriteByte, rewind

###

___fSum________________________________________________________________

Description: Use to report the 16-bit checksum of the specified

file.

Arguments:   The name of the file for which you want to report the

checksum.

Returns:     The checksum value for success, 0 for failure (int).

The return value is also placed in the z register.

Example:     a = fsum "filename"   /* Place the file's checksum */

/* value in the a register. */

print a               /* Display the checksum value */

/* in the WinFIRST report */

/* window. */

Also See:    fLength

###

___fwhm_______________________________________________________________

Description: Use to retrieve the peak width at half the absorbance

for the sample data peak nearest the supplied

wavenumber.

Arguments:   Three wavenumber values (int), where; the first value

is a wavenumber value for one end of the peak baseline,

the second value is the wavenumber used to locate the

peak, and the third value is the other end of the peak

baseline.

Returns:     The peak width at half the maximum absorbance, as

measured from the specified baseline (float).  The

return value is also placed in the z register.

Example:     load "pstyrene.abs"  /* Load file into memory. */

fwhm 2250            /* Scale to peak at 2250. */

###

___fwrite______________________________________________________________

Description: Use to write 8-byte double precision floating point

binary data to an open user file.  One value is written

at a time.

Arguments:   File number, from 1 to 16 (int ) and value to print to

file.

Returns:     1 for success, 0 for failure (int).

Example:     fopen 1 "file1.dat"  /* Open file1.dat for read and */

/* write. */

fwrite 1  2500.00    /* Write to the file. */

fclose 1             /* Close file number 1. */

Also See:    append, eof, exist, fclose, filepos, finput, fopen,

fprint, fread, freadBYte, fseek, fwriteByte, rewind

###

___fwriteByte__________________________________________________________

Description: Use to write raw ASCII data to a file.

Arguments:   File number, from 1 to 16 (int) and the registers that

contain values to write.

Returns:     Returns the number of bytes writen.  Also, places the

return value in the z register (int).

Example:     a = fwritebyte 1 a b c /* Write the values in a, b, */

/* and c to the current location in file 1. */

Also See:    append, eof, exist, fclose, filepos, finput, fopen,

fprint, fread, freadByte, fseek, fwrite, rewind

###

___getFirstDrive_______________________________________________________

Description: Use to return a string that contains the drive on which

WinFIRST is installed.  The returned

string takes the form: "c:", that is, a colon is

appended to the drive identifier.  This command is

useful for building directory names that contain data

for a macro, so that you can install the macro on

different drives and different computers.

Returns:     The drive desination for the drive on which WinFIRST or

software is installed (string).

Arguments:   None.

Example 1:   cd (getFirstDrive)  /* Change to the WinFIRST drive. */

Example 2:   string myDir

s = getFirstDrive

myDir = s + "\first\data\mydata"

cd myDir

###

___getFkey_____________________________________________________________

Description: Use to 1) pause with function keys displayed (as

defined with the fkeys command), 2) to display a

message in the status line below the displayed function

keys, and 3) to retrieve user input (that is, to

identify which function key the user selects).  Use the

getFkey command with the fkeys command to create

customized function keys for the WinFIRST macros you

write.

Arguments:   (optional)  Text of message to display on status line

while macro waits for user to choose a function key

(string).  If you do not provide message text, the

message "Press a Function Key..." appears on the status

line (string).

Returns:     The number of the function key the user seleted (int).

The return value is also placed in the z register.

Example:     int choice  /* Declare integer variable to hold */

/* key selection. */

int done    /* Declare integer variable for loop */

/* control. */

done = 0    /* Set done equal to 0. */

fkeys "Sample; Bkg; Plot; Done"

/* Define function keys.*/

while (done == 0)  /* Loop until done is non-zero. */

{                  /* Mark start of loop. */

choice = getFkey "Press a function key..."

/* Wait for user to press function key. */

/* The key user presses is stored in */

/* choice. */

if (choice == 1)   /* If first Fkey pressed... */

collectSample     /* Call this user function. */

else if (choice == 2) /* If second FKey pressed... */

collectBkg          /* Call this user function. */

else if (choice == 3) /* If third Fkey pressed... */

plotCurrentData     /* Call this user function. */

else if (choice == 4) /* If fourth Fkey pressed... */

done = 1 /* Set done to 1 to terminate loop. */

}            /* Mark end of the while loop. */

Also See:    fkeys, fkeysOff, menu

###

___getStringElement____________________________________________________

Description: Use to retrieve an individual character within a

string.

Arguments:   String from which to get the character (string) and

position of the character you want to retrieve (int).

Returns:     The ASCII value of the character (int).  To get the

actual character, assign the return value to a char

variable (as in Example 2 below).  The return value is

also placed in the z register.

Example 1:   x = getstringelement "This is a string" 4

/* The ASCII value of the 's' */

/* character (located in */

/* position 4) is stored in x. */

Example 2:   char ch             /* Declare a character variable. */

ch = getstringelement "This is a string" 4

/* The actual character 's' is */

/* stored in ch. */

Also See:    buildString, execString, isNumber, setStringElement,

string2Float, string2Int, stringCompare, stringIndex,

stringLeft, stringLength, stringLower, stringMid,

stringRight, stringUpper

###

___help________________________________________________________________

Description: Use to display information about a specific ApPro

command, or to display a list of all commands for

which online help available.  With this list displayed,

select a listed command to display help information

for that command.

NOTE: For best results, use a fixed-space font, such as

Courier, Fixedsys, or Terminal, to display ApPro help

text.  (ApPro help text is displayed using the

currently specified report font.  Choose Report from

the Display menu to specify the WinFIRST report font.)

A fixed-space font allows example code descriptions and

help text in general to display with the intended

spacing.

Arguments:   (optional)  Name of command for which you want to

display help information (string).  This string must be

enclosed by double quotation marks.

Example 1:   help getStringElement  /* Information for */

/* getstringelement is */

/* displayed. */

Example 2:   help  /* A list of available commands is displayed. */

/* Select a listed command to display help */

/* for that command. */

###

___if__________________________________________________________________

Description: Use to test the truth of an expression, and to then

continue according to the results.  If the expression

evaluates to true, the block following the "if"

statement is executed, otherwise it is not.  An "else"

statement can follow immediately after the "if" block.

In this case, when the "if" block is not executed, the

"else" block is executed.

Example 1:   if (A > 3)

begin            /* Start of "if" code block. */

load "tmp.abs" /* If A is more than 3, load file. */

end              /* End of "if" code block. */

/* Otherwise, do nothing. */

Example 2:   if (A > 3)

begin             /* Start of "if" code block. */

load "tmp.abs"  /* If A is more than 3, load the */

/* tmp.abs file. */

end               /* End of "if" code block. */

else

begin             /* Start of "else" code block. */

load "pure.abs" /* If A not more than 3, load the */

/* pure.abs file. */

end               /* End of "else" code block. */

Also See:    break, choice, choice and branch, decision loop, for,

while

Refer also to the expanded discussion of if statements,

if-else statements, and other control statements in

Chapter 2 of "Using WinFIRST Macros and ApPro."

###

___include_____________________________________________________________

Description: Use to execute an ApPro macro.  You can use the

include command from the ApPro command line or from

within a macro.  For example, use the include command

within a macro to include a second macro that contains

a set of user-defined functions you want to use.

NOTE: The include command does not clear the ApPro

symbol table.  Keep this in mind when using the include

command to use a user-defined function.

NOTE: The difference between the include command and

the run command is that the run command lets you "pass"

command line arguments to the macro you execute.  That

is, you can place specific values on the same line with

the run command.  It is the programmer's responsibility

to retrieve these arguments from the command line

argument structure before using run to call another

file which also uses command line arguments.  Command

line arguments with the run command are optional.

Arguments:   The name of the macro to execute (string).

Returns:     0 (int).  The return value is also placed in the z

register.

Example:     Create these macro files, "mac1.mac" and "mac2.mac",

as follows:

The mac1.mac file contains:

print "In mac1.mac... calling mac2.mac"

include "mac2.mac"        /* Include mac2.mac file. */

print "Back in mac1.mac"  /* This line is executed */

/* when mac2.mac is done */

/* executing. */

The mac2.mac file contains:

print "In mac2.mac..."

print "At last line.  Returning to mac1.mac..."

To run these macros:

At the ApPro command line type:  run "mac1.mac"

WinFIRST software users can also execute this macro

by working with the WinFIRST Macro Toolbox.

###

___increment___________________________________________________________

Description: Use to increase the value of a register by 1.  This

command also works to increment an integer that is the

last character in a file name.  (For example, tmp3

becomes tmp4.)

Arguments:   Register to increase (register).

Returns:     The new value is the specified register for success, or

0 for failure.

Example:     float array usrarray[5]  /* Declare a floating point */

/* array that can hold up */

/* to 5 elements. */

x = 1          /* Set value of register x to 1. */

while (x < 6)  /* Loop until x is greater than 6. */

begin          /* Mark start of loop. */

enter a "Please enter a wavenumber."

usrarray:data[x] = a  /* Assign the value stored */

/* in register a to the xth */

/* position in the float */

/* array named usrarray. */

increment x  /* Increment value of x by 1. */

end            /* Mark end of loop. */

Also See:    decrement

###

___info________________________________________________________________

Description: Use to display the data information report for the

current sample in a WinFIRST report window or in the

Gensis text display area.

Arguments:   None.

Returns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:     info  /* Print the sample data information report. */

###

___inp_________________________________________________________________

Description: Use to read a byte from the specified input port.

Arguments:   The input port (an integer ranging from 0 to 65535).

Returns:     The byte read from the port for success, 0 for failure

(int).  (A random value is placed in the z register.)

Example:     byte portinput          /* Declare a byte to hold */

/*port input. */

portinput = (inp 0x61)  /* Read in a byte from port */

/* 0x61 (hex). */

Also See:    inpw, outp, outpw

###

___inpw________________________________________________________________

Description: Use to read a word from the specified input port.

Arguments:   The input port (an integer ranging from 0 - 65535)

Returns:     The word read from the port for success, 0 for failure

(int).  (A random value is placed in the z register.)

Example:     int portinput             /* Declare an integer to */

/* hold port input. */

portinput = (inpw 0x61)]  /* Read in a word from */

/* port 0x61 (hex). */

Also See:    inp, outp, outpw

###

___integrate___________________________________________________________

Description: Use to find the area of a data region defined by two

specified wavenumbers and a calculated baseline that

connects the two wavenumber-defined data points.  The

integrate command fills in the data area that is

integrated, and displays an integration report in a

WinFIRST report window.

Arguments:   Two wavenumbers (float).

Returns:     The calculated value (float).  The return value is also

placed in the z register.

Example:     load "pstyrene.abs"   /* Load file into memory. */

abc                   /* Correct baseline */

/* automatically. */

integrate  3000 2500  /* Find area between */

/* 3000 and 2500 wavenumbers. */

###

___interpolate_________________________________________________________

Description: Use to estimate a different digital resolution.

Arguments:   The new resolution, from 0.1 to 64.0 wavenumbers

(float).

NOTE: If the value is out of range, no estimate is

made.

Returns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:     load "pstyrene.abs"  /* Load the file into memory. */

interpolate 4.0      /* Change data resolution to */

/* 4.0 wavenumbers. */

show                 /* Display the resulting data. */

###

___invert______________________________________________________________

Description: Use to invert (flip) the display of sample data.  For

example, use this command when you want to find peaks

in single beam data.

Arguments:   None.

Returns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:     load "pstyrene.abs"  /* Load the file into memory. */

invert               /* Invert the sample data. */

show                 /* Display the resulting data. */

###

___irChart_____________________________________________________________

Description: Use to display the WinFIRST Correlation Chart workset.

Arguments:   (optional)  The name of the correlation chart to use.

If you do not provide a name, the name of the last

correlation chart you worked with is used.

Returns:     0 (int).

Example:     load "pstyrene.abs"  /* Load file into memory. */

irChart charta       /* Display the correlation */

/* chart workset and load */

/* the charta.cht */

/* correlation chart. */

###

___isNumber____________________________________________________________

Description: Use to determine whether the specified value is a

valid number.

Arguments:   Any literal value, variable, or register.

Returns:     1 for success, 0 for failure (int).

Example:     fopen 1 "testfile.dat"  /* Open testfile.dat file. */

finput 1 a              /* Read first line in file. */

fclose 1                /* Close the file. */

if (isnumber a)  /* If text in a is valid number... */

print "The value stored in a is a valid number."

/* Display the message. */

Also See:    buildString, execString, getStringElement,

setStringElement, string2Float, string2Int,

stringCompare, stringIndex, stringLeft, stringLength,

stringLower, stringMid, stringRight, stringUpper

###

___kk__________________________________________________________________

Description: Use to perform a Kramers�Kronig transformation on a

transmittance spectrum.

Arguments:   None.  Transform is performed on current sample.

Returns:     1 for success, 0 for failure (int).

Example:     kk

###

___launch______________________________________________________________

Description: Use to execute a Windows or DOS application

concurrently with WinFIRST.  (WinFIRST does not need

to wait for the application to terminate.)

If the application is unsuccessful in starting,

ApPro displays an error message.

Arguments:   The command that starts the application (string).

Usually this consists of a program name and 0 or more

command line arguments.

Returns:     1 for success, 0 for failure (int).  (A random value is

placed in the z register.)

Example:     launch "pbrush.exe"

Also See:    dos, execute, exit, shell, system

###

___library_____________________________________________________________

Description: Library is an object with prperties and methods you

can use to create and modify user libraries.

NOTE: The library object works with user libraries only.

See examples for using the library object at the end of

this description.

Properties:  Properties for the library object are:

name:          Use to specify the user library to use, that is,

the name of library that is to be the current

library.

Arguments:     The name of the user library to use (string).

Syntax:        library:name = "lib1"

storeResults:  Use before a library search 1) to place in the x

register the name of the library spectrum that is

the top hit and 2) to place in the y register the

library spectrum's index number.

Arguments:     1 to store the library search results or 0 to

disable the storeResults method.

Syntax:        library:storeResults=1

Methods:     Methods for the library object are:

add:           Use to add the current sample data to the current

library.  The data is always added to the end of

the library.  Returns 0 for success.  Returns 1

for failure.

Arguments:     None.

Syntax:        library:add

numSpectra:    Use to return the number of spectra in the

current library.  Retruns 0 if the library does

not exist or if no spectra are in the library.

Arguments:     None.

Syntax:        library:numspectra

retrieve:      Use to retrieve the spectrum of the specified

index from the current library, and place that

spectrum in a new data window.  Returns 0 if

successful.  Returns 1 for failure.

Arguments:     The library index for the spectrum to retrieve.

Syntax:        library:retrive index

Example 1:     /* This macro plots the entire contents of a */

/* library.  It assumes a plot method named */

/* "libplot" has already been created using */

/* the plot composer. */

library:name = "icon-1"

a = library:numSpectra

if (ask "Prints " a " pages.  Press N to cancel.")

{

for i = 1 to a step 1

{

if (library:retrieve i)

{

plot "libplot"

/* libplot is a plot method. */

sample:clear  /* So 500 data windows */

/* are not opened. */

}

else

{

print "ERROR getting " i " from icon-1"

break;

}

}

}

Example 2:     /* This macro puts the titles of all spectra */

/* in a library into a report window.  Note the */

/* use of the setStringElement command to get a */

/* tab stop into a string.  The tab character */

/* ASCII 9. */

string tab

tab = " "

tab = setStringElement tab 1 9

/* Above changes the space to a tab. */

library:name = "icon-1"

a = library:numSpectra

report:new

for i = 1 to a step 1

{

if (library:retrieve i)

{

s = buildString i tab sample:name tab sample:title

print s

sample:clear

}

else

{

print "ERROR getting " i " from icon-1"

break

}

}

Also See:    add2Lib

###

___lineout_____________________________________________________________

Description: Use to replace the data between two specified

wavenumbers with a straight line.

Arguments:   The starting and ending wavenumbers of the data you

want to line out.

Returns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:     load "pstyrene.abs"  /* Load file into memory. */

show                 /* Display the spectrum. */

lineout 4000 3200    /* Replace the data between */

/* 4000 and 3200 wavenumbers */

/* with a straight line. */

show                 /* Display the results. */

###

___linreg______________________________________________________________

Description: Linear regression.  Use to calculate the slope,

intercept, and standard deviation of peaks in the

current sample.  Results are displayed in a WinFIRST

report window.

Arguments:   None.

Returns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:     load "pstyrene.abs"  /* Load file into memory. */

autoscale            /* Display the data autoscaled. */

abc                  /* Correct the baseline */

/* automatically. */

linreg               /* Calculate the linear */

/* regression of the peaks. */

###

___ln__________________________________________________________________

Description: Use to calculate the natural logarithm of the given

value, or of the current sample if no value is given.

Arguments:   (optional)  Value for which you want to compute the

natural log (float).

Returns:     If no value is specified, 1 (int) is returned and is

placed in the z register.  If a value is specified, the

result (float) is returned, and the z register is not

modified.

Example 1:   load "pstyrene.abs"  /* Load file into memory. */

ln                   /* Calculate ln of each */

/* data point. */

show                 /* Display the resulting data. */

Example 2:   print (exp 2.5)      /* Calculate ln of 2.5. */

Also See:    exp, log

###

___load________________________________________________________________

Description: Use to open (or load) a data file into memory

and display the data in the data display area.

Arguments:   (optional for WinFIRST only)  The name of the file you

want to load.  If the file is not in the current

directory, ApPro searches all other WinFIRST data

directories for the file.  If the file is not in a

WiFIRST directory, you must specify the complete

file path (including drive designation).  If you

provide no argument, a multiple-file open dialog

appears.

Returns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:     load "c:\winfirst\data\galaxy\pure1.abs"

/* Load the file into memory*/

/* and display it in the data */

/* display area. */

Also See:    reload, sample, show

###

___loadRef_____________________________________________________________

Description: Use to open (or load) one or multiple data files

into the active data window as reference data

files.  (WinFIRST allows up to 32 reference data

files in each data window.)

Arguments:   (optional)  The name of the file you want to load.

If the file is not in the current directory, ApPro

searches all other WinFIRST data directories for the

file.  If the file is not in a WinFIRST directory, you

must specify the complete file path (including drive

designation).  If you provide no argument, a multiple-

file open dialog appears.

Retruns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:     loadRef "c:\winfirst\data\galaxy\pure1.abs"

/* Load the file into memory and display */

/* it as reference data in the active */

/* data window. */

Also See:    reference

###

___log_________________________________________________________________

Description: Use to calculate the base 10 logarithm of the given

value, or of the current sample if no value is given.

Arguments:   (optional)  Value for which you want to compute the

base 10 log (float).

Returns:     If no value is specified, 1 (int) is returned and is

placed in the z register.  If a value is specified, the

result (float) is returned, and the z register is not

modified.

Example 1:   load "pstyrene.abs"    /* Load file into memory. */

log                    /* Calculate log of each */

/* sample file data point. */

show                   /* Display resulting data. */

Example 2:   print (log 2.5)        /* Calculate the log of 2.5. */

Also See:    exp, ln

###

___math expressions____________________________________________________

Description: You can use various ApPro math expressions.  However,

you should have a good working knowledge of ApPro math

commands before using this advanced feature.  See

Chapter 2 in "Using WinFIRST Macros and ApPro" for

more information.

###

___matt2dx_____________________________________________________________

Description: Use to convert a copy of the current sample data

to the JCAMP file format.  You can convert four data

types: absorbance, transmittance, diffuse reflectance,

and single beam.  The original data file is not

affected.

Arguments:   (optional)  The name of the converted file (string).

if you do not specify a name for the converted file,

the sample file filename is used with .dx assigned

as the filename extension.

Returns:     1 for success, 0 for failure (int).  Also, the return

value is stored in the z register.

Example:     load "pstyrene.abs"  /* Load file into memory. */

matt2dx              /* Convert file to the JCAMP */

/* format and use pstyrene.dx */

/* as the file name. */

matt2dx "pstyren2"   /* Convert file to the JCAMP */

/* format and use pstyren2.dx */

/* as the file name. */

Also See:    dx2matt, matt2nicDX, matt2nicSX, matt2omnic, matt2pcir,

matt2sadtler, matt2spc, matt2xy, nic2matt,

sadtler2matt, spc2matt, xy2matt

###

___matt2nicDX__________________________________________________________

Description: Use to convert the current sample data to the Nicolet

DX data file format.  The filename extension ".nic"

is assigned to the converted file.

Arguments:   (optional)  The name of the converted file (string).

If you do not specify a name for the converted file,

the sample file filename is used (with .nic as the

filename extension).

Returns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:     load "pstyrene.abs"  /* Load the file into memory */

matt2nicDX           /* Convert file to Nicolet DX */

/* format and use pstyrene.nic */

/* as the file name. */

Also See:    dx2matt, matt2dx, matt2nicSX, matt2omnic, matt2pcir,

matt2sadtler, matt2spc, matt2xy, nic2matt,

sadtler2matt, spc2matt, xy2matt

###

___matt2nicSX__________________________________________________________

Description: Use to convert the current sample data to the Nicolet

SX data file format.  The filename extension ".nic"

is assigned to the converted file.

Arguments:   (optional)  The name of the converted file (string).

If you do not specify a name for the converted file,

the sample file filename is used (with .nic as the

filename extension).

Returns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:     load "pstyrene.abs"   /* Load the file into memory */

matt2nicSX    /* Convert file to Nicolet SX format */

/* and use pstyrene.nic as the name. */

Also See:    dx2matt, matt2dx, matt2nicDX, matt2omnic, matt2pcir,

matt2sadtler, matt2spc, matt2xy, nic2matt,

sadtler2matt, spc2matt, xy2matt

###

___matt2omnic__________________________________________________________

Description: Use to convert the current sample data to the Nicolet

omnic data file format.  The filename extension ".nic"

is assigned to the converted file.

Arguments:   (optional)  The name of the converted file (string).

If you do not specify a name for the converted file,

the sample file filename is used (with .nic as the

filename extension).

Returns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:     load "pstyrene.abs"  /* Load the file into memory */

matt2omnic           /* Convert file to Nicolet omnic */

/* format and use pstyrene.nic */

/* as the file name. */

Also See:    dx2matt, matt2dx, matt2nicDX, matt2nicSX, matt2pcir,

matt2sadtler, matt2spc, matt2xy, nic2matt,

sadtler2matt, spc2matt, xy2matt

###

___matt2pcir___________________________________________________________

Description: Use to convert the current sample data to the Nicolet

pcir data file format.  The filename extension ".nic"

is assigned to the converted file.

Arguments:   (optional)  The name of the converted file (string).

If you do not specify a name for the converted file,

the sample file filename is used (with .nic as the

filename extension).

Returns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:     load "pstyrene.abs"   /* Load the file into memory */

matt2pcir     /* Convert file to Nicolet pcir format */

/* and use pstyrene.nic as file name. */

Also See:    dx2matt, matt2dx, matt2nicDX, matt2nicSX, matt2omnic,

matt2sadtler, matt2spc, matt2xy, nic2matt,

sadtler2matt, spc2matt, xy2matt

###

___matt2sadtler________________________________________________________

Description: Use to convert the current sample data to the Sadtler

data file format (.irf).

Arguments:   (optional)  The name of the converted file (string).

If you do not specify a name for the converted file,

the sample file filename is used with .irf as the

filename extension.

Returns:     1 for success, 0 for failure (int).  The return value is

stored in the z register.

Example:     load "pstyrene.abs"  /* Load the file into memory */

matt2sadtler         /* Convert file to the Sadtler */

/* format and use pstyrene.irf */

/* as the file name. */

matt2sadtler "pstyr2"  /* Convert file to the */

/* Sadtler format and use */

/* pstyr2.irf as a file name. */

Also See:    dx2matt, matt2dx, matt2nicDX, matt2nicSX, matt2omnic,

matt2pcir, matt2spc, matt2xy, nic2matt, sadtler2matt,

spc2matt, xy2matt

###

___matt2spc____________________________________________________________

Description: Use to convert the current sample data to the Spectra

Calc file format.

Arguments:   (optional)  The name of the converted file (string).

If you do not specify a name for the converted file,

the sample file filename is used with .spc assigned as

the filename extension.

Returns:     1 for success, 0 for failure (int).  Also, the return

value is stored in the z register.

Example:     load "pstyrene.abs"    /* Load the file into memory */

matt2spc     /* Convert file to the Spectra Calc */

/* format and use pstyrene.spc as the */

/* file name. */

matt2dx "pstyren2"

/* Convert file to the Spectra Calc */

/* format and use pstyren2.spc as the */

/* as the file name. */

###

___matt2xy_____________________________________________________________

Description: Use to write the current sample data to a file as a

table of x and y values.

Arguments:   (optional)  The name of the converted file (string).

If you do not specify a name for the converted file,

the sample file filename is used with .tbl assigned as

the filename extension.

Returns:     1 for success, 0 for failure (int).  Also, the return

value is stored in the z register.

Example:     load "pstyrene.abs"  /* Load file into memory. */

matt2xy              /* Convert file to an xy table */

/* and use pstyrene.tbl as the */

/* file name. */

matt2xy "pstyren2"   /* Convert file to an xy table */

/* and use pstyren2.tbl as the */

/* file name. */

###

___max_________________________________________________________________

Description: Use to find the maximum y axis value for the current

sample data.

Arguments:   None.

Returns:     The y maximum (float).  The return value is also

placed in the y register.

Example:     load "pstyrene.abs"  /* Load file into memory. */

max                  /* Retrieve largest y value. */

###

___maxX________________________________________________________________

Description: Use to find the maximum x axis value for the current

sample data.

Arguments:   None.

Returns:     The x maximum (float).  The return value is also

placed in the z register.

Example:     load "pstyrene.abs"  /* Load file into memory. */

maxx                 /* Retrieve largest x value. */

###

___md__________________________________________________________________

Description: Use to create a new DOS directory.

Arguments:   The name of the directory to create.

Returns:     1 (int).  The return value is also placed in the z

register.

Example:     md c:\first\data\mine  /* Create the directory. */

/* MD is a DOS command. */

md junk

Also See:    cd, copy, dir, erase, rename

###

___menu________________________________________________________________

Description: Use to place a menu box on the screen.  Once the menu

box is displayed, select any numbered menu item in one

of three ways: 1) with the mouse, point to the item and

click the left mouse button, 2) press the number

associated with the menu item, or 3) use the arrow keys

to place the highlight bar over the item and then press

the Enter key to select the highlighted item.

Arguments:   The list of the text items to be placed in the menu

(string).  The list is one text string that contains

each menu text item separated by a semi-colon (;).  For

example: "Sample;Background;Plot;Quit"

Returns:     The number of the menu item that is selected (int).  If

the Escape key is pressed, or if the mouse is clicked

outside the menu box, negative one (-1) is returned (the

menu has been terminated without a choice).  The return

value is also placed in the z register.

Example:     int choice  /* Declare integer variable to hold */

/* item selection. */

int done    /* Declare integer variable for loop */

/* control. */

done = 0    /* Set done equal to 0. */

while (done == 0)  /* Loop until done is non-zero. */

{           /* Mark start of loop. */

choice = menu "Sample;Bkg;Plot;Quit"

/* Wait for menu item to be selected, */

/* and after item is selected store the */

/* item number in choice. */

if (choice == 1)  /* If first item selected... */

collectSample   /* Call this user function. */

else if (choice == 2)/* If second item selected... */

collectBkg         /* Call this user function. */

else if (choice == 3)/* If third item selected... */

plotCurrentData    /* Call this user function. */

else if (choice == 4)/* If fourth item selected... */

done = 1 /* Set done to 1 to terminate the loop. */

else if (choice == -1) /* If menu is terminated... */

done = 1 /* Set done to 1 to terminate the loop. */

}            /* Mark the end of the while loop. */

Also See:    fkeys, fkeysOff, getFkey

###

___min_________________________________________________________________

Description: Use to find the minimum y axis value for the current

sample.

Arguments:   None.

Returns:     The y minimum for success, 0 for failure (float).

The return value is also placed in the y register.

Example:     load "pstyrene.abs"  /* Load file into memory. */

min                  /* Retrieve lowest y value. */

Also See:    minx

###

___minX________________________________________________________________

Description: Use to find the minimum x axis value for the current

sample.

Arguments:   None.

Returns:     The x minimum for success, 0 for failure (float).  The

return value is also plaed in the z register.

Example:     load "pstyrene.abs"  /* Load file into memory. */

minx                 /* Retrieve lowest x value. */

Also See:    min

###

___multiply____________________________________________________________

Description: Use to multiply each data point of the current sample

with the corresponding data of another file.   The

resulting data becomes the sample data, and the

original sample data becomes the R1 reference data

file.

Arguments:   The name of the data file by which you want to multiply

the sample data in memory.

Returns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:     load "pure1.abs"       /* Load file into memory. */

multiply "pure2.abs"   /* Multiply pure1.abs by the */

/* data in pure2.abs file. */

title "pure1 * pure2"  /* Title the results. */

show                   /* Display the results. */

Also See:    add, average, divide, subtract

###

___neta________________________________________________________________

Description: Use to calculate the net absorbance for the specified

wavenumber to the specified baseline.

Arguments:   Three wavenumbers (float).  The two extreme wavenumbers

define the baseline.  The wavenumber between the two

extremes is the data point for which net absorbance is

calculated.

Returns:     The net absorbance for success, 0 for failure (float).

The return value is also placed in the z register.

Example:     load "pstyrene.abs"  /* Load file into memory. */

neta 1500.0 1600.0 1550

/* Calculate net absorbance */

/* from the data point at */

/* 1550 wavenumbers to the */

/* baseline defined by 1500 */

/* and 1600 wavenumbers. */

###

___nic2matt____________________________________________________________

Description: Use to a translate a copy of a file in a Nicolet format

(dx, pcir, sx, or omnic) to a file in the Mattson

FIRST data file format.  The translated file becomes

the current sample.  Four data types (stored in the

Nicolet format) can be translated: absorbance,

transmittance, diffuse reflectance, and single beam.

An appropriate Mattson filename extension is

assigned to the translated file.  The original Nicolet

file that is translated is not removed or otherwise

altered by the nic2matt command.  You can use the

translated data file in the same way you would any data

file created by WinFIRST.

Arguments:   The directory path and name of Nicolet data file

(string).

Returns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:     nic2matt "\nic\data\pstyrene"

/* Translate file and make it the */

/* sample data file. */

Also See:    dx2matt, matt2dx, matt2nicDX, matt2nicSX, matt2omnic,

matt2pcir, matt2sadtler, matt2spc, matt2xy,

sadtler2matt, spc2matt, xy2matt

###

___normalize___________________________________________________________

Description: Use to move the lowest sample data data point to zero

and also move all other data points by the same amount.

Arguments:   None.

Returns:     1 (int).  The return value is also placed in the z

register.

Example:     load "pure2.abs"  /* Load file into memory. */

autoscale         /* Display auto scaled data. */

normalize         /* Normalize the data. */

show              /* Display the results. */

###

___offset______________________________________________________________

Description: Use to offset the data along the y axis by a specified

value.

Arguments:   The offset value (float).

Returns:     1 (int).  The return value is also placed in the z

register.

Example:     load "pure2.abs"  /* Load file into memory. */

autoscale         /* Display autoscaled data. */

offset 2.0        /* Offset by 2.0 absorbance units. */

show              /* Display the results. */

###

___offsetReferences____________________________________________________

Description: Use to offset all compatible reference data files along

the y axis to a sample data value you provide as an

argument or choose with the mouse.

NOTE: A reference data file is compatible with the

sample data when both files interferograms, both files

are single beam data or both files are a spectral data

type (.ras, .abs, or .drt).

Reference files that are not compatible with the sample

data are not affected by the offset procedure.

Arguments:   (optional)  An x-axis value on the sample data from

which the y-axis value is taken (float).

NOTE: If no argument is provided, the mouse pointer

cross hairs snaps to the sample data and a dialog

appears to let you choose the offset sample data value.

Example:     load "m-xyl.abs"      /* Load file into memory. */

autoscale             /* Display autoscaled data. */

loadref "p-xyl.abs"   /* Load file into memory. */

offsetreferences 3200 /* Offset references to the */

/* specified sample data */

/* value. */

show                  /* Display the results. */

###

___outp________________________________________________________________

Description: Use to write a byte to the specified output port.

Arguments:   The port to which to send the byte (integer between

0 and 65535), and the value to write to the port

(integer between 0 and 255).

Returns:     The value that is written to the port (byte).

WARNING:     This command is potentially very dangerous!  Before

you write to a port, make sure you know what the port

controls and that you understand the consequences of

your actions.

Example:     outp 0x61  182  /* Write 182 to port 0x61 (hex). */

Also See:    inp, inpw, outpw

###

___outpw_______________________________________________________________

Description: Use to write a word to the specified output port.

Arguments:   The port to which to send the word (integer between 0

and 65535), and the value to write to the port

(integer between 0 and 65535).

Returns:     The value that is written to the port (byte).

WARNING:     This command is potentially very dangerous!  Before

you write to a port, make sure you know what the port

controls and that you understand the consequences of

your actions.

Example:     outpw 0x61  1232  /* Write 1232 to port 0x61 (hex). */

Also See:    inp, inpw, outp

###

___overlay_____________________________________________________________

Description: Use to overlay the sample and reference spectra.  Also

use to load a reference spectrum to overlay with the

sample (and other loaded references).

Arguments:   (optional)  The name of the reference data file to load

and overlay.  If you specify no file name, the current

references and sample are overlaid, one on top of

another.

Returns:     0 (int).  The return value is also placed in the z

register.

Example:     load "mix2.abs"    /* Load file into memory. */

overlay "mix.abs"  /* Overlay mix.abs with mis2.bas */

/* and display the results. */

Also See:    stack, superimpose

###

___p2pNoise____________________________________________________________

Description: Use to calculate peak-to-peak noise for the current

sample spectrum within a wavenumber range you specify.

If verbose is set to "on" the calculated value is

automatically printed to the screen.

Arguments:   The starting and ending wavenumbers (float).

Returns:     The calculated value for success, 0 for failure

(float).  Also the return value is stored in the z

register.

Example:     load "pstyrene.abs"  /* Load file into memory. */

p2pnoise 2000 2250   /* Calculate peak-to-peak noise */

/* between 2000 and 2250 cm-1. */

print z              /* Print the results. */

###

___parms_______________________________________________________________

Description: Parms is an object that contains properties and

methods you set and use to control several aspects of

the WinFIRST software environment.  All properties and

methods for the parms object are described below.

NOTE: New settings for properties and methods you

specify for the parms object take place only after the

displayed sample data is redisplayed, using, for

example, the show or expand commands.

Properties:  You can get the current setting for any property

specified by parms.  For example:

r = parms:bkgcolor  /* Store the current setting for */

/* the bkgcolor property in the */

/* r register. */

You can also set a new setting for any property

specified by parms.  For example:

parms:labelshadowon = 1  /* Set the setting for the */

/* labelshadowon property */

/* to 1. */

For the first example, the return value matches the

value of the property setting itself (integer, string,

etc.).  For the second example, the value you specify

must also match the value of the property setting.

Each property value is included with the property

description below.

NOTE: For properties with a Color argument, the color

is represented by a 24-bit integer (which is easy to

represent with hexadecimal numbers).  The color is

determined by setting percentages of blue, green, and

red.  For instance, 0xFF00FF means 100% blue, 0% green,

and 100% red.  The leftmost FF represents percentage of

blue, the center values (00) represent percentage of

green, and the rightmost FF represents percentage

of red.  So, 0xFF0000 = 100% blue only, 0x00FF00 = 100%

green only, and 0x0000FF = 100% red only.

Properties for the parms object are:

auditLevel:    Use to turn audit trail use on and off.

Arguments:     0 = Off, 1 = On (int).

Syntax:        parms:auditlevel = 0

axesColor:     Use to set the axes color.

Arguments:     Color (int).

Syntax:        parms:axisColor = 0x00FF00

bkgColor:      Use to set the background color for various

items in the software.

Arguments:     Color (int).

Syntax:        parms:bkgcolor = 0x00FF00

bottom:        Use to specify the minimum y-axis value to display

(when specified scaling limits are in use.)

Arguments:     The y-axis value (float).

Syntax:        parms:bottom = 0.5

chart:         Use to set the name of the current functional

group correlation chart.

Arguments:     Name of chart (string).

Syntax:        parms:chart = "charta"

closePrompt:   Use to specify when WinFIRST should prompt you to

save data.

Arguments:     The name of the prompt setting you want to use

(string): Always, Never, or Sometimes (that is,

only when  data has changed).

Syntax:        parms:closeprompt = "always"

company:       Use to set the company name used by the WinFIRST

log book.

Arguments:     Any name (string).

Syntax:        parms:company = "Mattson"

cwd:           Use only to report (not to set) the current

working directory.

Arguments:     None.

Syntax:        x = parms:cwd;print x

ddeTimeOut:    Use to specify the time, in seconds, in which a

dde command must be acknowledged or a timeout

error (and therefore an execution error) will

occur.

NOTE: Formerly, this value was hard-coded

internally to three (3) seconds.  Now (in WinFIRST

version 3.0 and above), the default time is 10

seconds, and you can change this setting using the

parm object's ddeTimeOut property.

Arguments:     The number of seconds to wait before a timeout

error occurs.

Syntax:        parms:ddeTimeOut = 5

defaultDataType: Use to specify the default data type to load.

Arguments:     The name of the data type to load (string):

Absorbance, Transmittance, Diffuse Reflectance,

Single Beam, Interferogram, JCAMP-DX, ASCII Table,

Spectra-Calc, and Sadtler.

Syntax:        parms:defaultdatatype = "interferogram"

displayLimits: Use to set the data scaling method.

Arguments:     The name of the scaling method to use (string):

autoscale, current (as specified using the left,

right, top and bottom properties), or defaults.

Syntax:        parms:displaylimits = "autoscale"

displaymethod: Use to specify the name of a display template to

specify how new data and report windows you open

are displayed in the WinFIRST workspace.

Arguments:     The name of the display template to use (string)

or <None> to use no display template (string).

Syntax:        parms:displaymethod = "example.ini"

parms:displaymethod = "<None>"

displaymethodautoscale:  Use to turn use of the Autoscale to

Screen global option (Display Composer) on and

off.

Arguments:     Either 1 or 0 to indicate on or off, respectively.

Syntax:        parms:displaymethodautoscale = 1

displaymethodrecycle:  Use to turn use of the Recycle WIndows

global option (Display Composer) on and off.

Arguments:     Either 1 or 0 to indicate on or off, respectively.

Syntax:        parms:displaymethodrecycle = 0

displayMode:   Use to set the data display mode.

Arguments:     The name of the display mode type you want to use

(string): Linear, Expanded, or Microns.

Syntax:        parms:displaymode = "linear"

enableCtrlBrk: Use to control whether a user is allowed to halt

a macro's execution by pressing the Control +

Break key combination.  The default value is 1, to

indicate checking for the Control + Break key

combination is on.

Arguments:     Either 1 or 0 to indicate the status of Control +

Break checking, where 1 indicates checking is in

use and 0 indicates the checking is not is use (so

that a macro's execution can not be halted with

Control + Break).

Syntax:        parms:enableCtrlBrk = 0

enableUndo:    Use to enable the undo option.

Arguments:     The name of the setting to use (string): on or off.

Syntax:        parms:enableUndo = "on"

fileOverwritePrompt: Use to enable the file overwrite prompt.

Arguments:     The name of the setting to use (string): on or

off.

Syntax:        parms:fileoverwriteprompt = "on"

fileType:      Use to specify the default data type to load.

Arguments:     The name of the data type to load (string):

Absorbance, Transmittance, Diffuse Reflectance,

Single Beam, Interferogram, JCAMP-DX, ASCII Table,

Spectra-Calc, and Sadtler.

Syntax:        parms:filetype = "interferogram"

grid:          Use to enable display of a line grid.

Arguments:     The name of the setting to use (string): on or

off.

Syntax:        parms:grid = "on"

gridColor:     Use to set the grid color.

Arguments:     Color (int).

holeColor:     Use to set the color for the data display area

background.

Arguments:     Color (int).

Syntax:        parms:holecolor = 0x00FF00

labelColor:    Use to set the axes and axes label color.

Arguments:     Color (int).

Syntax:        parms:labelcolor = 0x00FF00

left:          Use to set the leftmost wavenumber to display

(when specified scaling limits are in use).

Arguments:     The wavenumber (float).

Syntax:        parms:left = 3000.0

lockouts:      Use to specify whether you want to lock out all

WinFIRST procedures that change program settings.

Arguments:     0 = Off, 1 = On (int).

Syntax:        parms:lockouts = 0

mathMethod:    Use specify the current math method (math

configuration).

Arguments:     Any valid math method file name (string).

Syntax:        parms:mathmethod = "special"

maxNewWindows: Use to specify whether new data windows are

maximized.

Arguments:     The name of the setting to use: on or off.

Syntax:        parms:maxnewwindows = "off"

monochrome:    Use to specify monochrome or color display.

Arguments:     1 for monochrome, 0 for color (int).

Syntax:        monochrome = 0

name:          Use to specify the base name for the current set of

parameter files.

Arguments:     The name of the file (string).

Syntax:        parms:name = "special"

operator:      Use to specify the name of the current operator.

This name is otherwise supplied as the operator

logs on as the WinFIRST software starts

(when operator log on is in use).  Also, it is the

current operator name that appears in the sample

data tagger (when taggers are in use and the

tagger is defined to include the operator name).

Arguments:     Operator name (string).

Syntax:        parms:operator = "Sean Ryan Callaghan"

overlayStyle:  Use to set the data display style.

Arguments:     1, 2, 4, or 8 (int), where 1 = Stack, 2 = Overlay,

4 = Superimpose, and 8 = Sample Only.

plotMethod:    Use specify or report the current plot method

(plot configuration).

Arguments:     Any valid plot method file name (string).

Syntax:        parms:plotmethod = "special"

ref1Color:     Use to set the R1 (first) reference color.

Arguments:     Color (int).

Syntax:        parms:ref1color = 0x00FF00

ref2Color:     Use to set the R2 (second) reference color.

Arguments:     Color (int).

Syntax:        parms:ref2color = 0x00FF00

ref3Color:     Use to set the R3 (third) reference color.

Arguments:     Color (int).

Syntax:        parms:ref3color = 0x00FF00

ref4Color:     Use to set the R4 (fourth) reference color.

Arguments:     Color (int).

Syntax:        parms:ref4color = 0x00FF00

refTitle:      Use to specify whether reference file titles are

displayed.

Arguments:     The setting you want to use (string): on or off.

Syntax:        parms:reftitle = "off"

reportColor:   Use to set the color for report text.

Arguments:     Color (int).

Syntax:        parms:reportcolor = 0x00FF00

right:         Use to set the rightmost wavenumber displayed

(when specified scaling limits are in use).

Arguments:     The wavenumber (float).

Syntax:        parms:right = 500

rptBkgColor:   Use to set the color for the report area

background.

Arguments:     Color (int).

Syntax:        parms:rptbkgcolor = 0x00FF00

sampleColor:   Use to change the sample data color.

Arguments:     Color (int).

scaling:       Use to set the data scaling method.

Arguments:     The number that represents the scaling method

(int), where 1 = autoscale, and 2 = use set

scaling limits.  (These limits are set using the

left, right, top and bottom parameters.)

Syntax:        parms:scaling = 2

scanMethod:    Use to specify the name of the data collection

method to use as the current method.

Arguments:     Method name (string).

Syntax:        parms:scanmethod = "special"

signOn:        Use to specify whether you want to display the

WinFIRST Log Book dialog each time you start

WinFIRST and to also require entry of a log book

record.  If the operator chooses Cancel from the

Log Book dialog to avoid entering a log book

record, WinFIRST closes.

Arguments:     0 = Off, 1 = On (int).

Syntax:        parms:signon = 1

simpleMenus:   Use to specify whether you want WinFIRST to start

in simple menus mode.

Arguments:     0 = No, 1 = Yes (int).

Syntax:        parms:simplemenus = 0

smooth_method: Use to set the smooth method to use for spectral

smoothing.

Arguments:     The number that represents the smooth methods

(int), where 0 = Boxcar, 1 = Cubic, 2 = Quartic,

3 = Quadratic, and 4 = Quintic.

Syntax:        parms:smooth_method = 0

smooth_points: Use to specify the number of smoothing points to

use for spectral smoothing.

Arguments:     Any odd number from 3 to 25 (int).

Syntax:        parms:smooth_points = 7

stackSize:     Use to set the number of displayed data files.

(One sample file and up to 32 reference files.)

Arguments:     A number from 1 to 33 to specify the total number

of data files to display (int).

Syntax:        parms:stacksize = 4

tabColor:      Use to set the name color for the inactive tab

keys.

Arguments:     Color (int).

Syntax:        parms:tabcolor = 0x00FF00

tagger:        Use to specify whether a sample data tagger is

used.

Arguments:     1 = Yes, 0 = No (int).

tagString:     Use to define the current tagger string.

Arguments:     Tagger string (string).

NOTE: Refer also to Chapter 7 of the "WinFIRST

User's Reference Guide."

Syntax:        parms:tagstring = "%1^n%2^n%3

testLevel:     Use to set the scan test level.

Arguments:     The number that represents the test level (int),

where 0 = standard test, 1 = interferogram test,

2 = linearity and water test, and 3 = 100% line

noise test.

Syntax:        parms:testlevel = 2

title:         Use to specify whether the sample file title is

displayed.

Arguments:     The setting you want to use (string): on or off.

Syntax:        parms:title = "off"

titleColor:    Use to specify the color used to display the

sample file title.

Arguments:     Color (int).

Syntax:        parms:titlecolor = 0x00FF00

top:           Use to set the maximum y-axis value to display

(when specified scaling limits are in use).

Arguments:     The y-axis value (float).

Syntax:        parms:top = 2.5

xPrecision:    Use to set the number of decimal places to include

in reported x-axis values.

Arguments:     0 to 9 (int), where: 0 specifies no decimal places

and 9 specifies 9 decimal places.

Syntax:        parms:xprecision = 4

yPrecision:    Use to set the number of decimal places to include

in reported y-axis values.  (Also specifies the

number of decimal places to report with search

results.  However at leat two decimal places are

reported regardless of this setting.

Arguments:     0 to 9 (int), where: 0 specifies no decimal places

and 9 specifies 9 decimal places.

Syntax:        parms:yprecision = 2

Methods:     Methods for the parms object are:

read:          Use to load the current parms file.  (Get or set

the name of the current parms file with the name

property for the parms object.)

Arguments:     None.

Syntax:        parms:read

readDisplayCfg: Use to load the display configuration file you

identify as an argument.

Arguments:     The name of the display configuration file to

read (load).  Inclusion of the filename extension

(.ini) is optional.

Syntax:        parms:readdisplaycfg "usedlast"

Returns:       1 for success, 0 for failure (int).

readSearchCfg: Use to load the search configuration file you

identify as an argument.

Arguments:     The name of the search configuration file to read

(load).  Inclusion of the filename extension

(.ini) is optional.

Syntax:        parms:readsearchcfg "usedlast"

Returns:       1 for success, 0 for failure (int).

readMathCfg:   Use to load the math configuration file you

identify as an argument.

Arguments:     The name of the math configuration file you want

to read (load).  Inclusion of the filename

extension (.ini) is optional.

Syntax:        parms:readmathcfg "usedlast"

Returns:       1 for success, 0 for failure (int).

write:         Use to save the current parms file.  (Get or set

the name of the current parms file with the name

property for the parms object.)

Arguments:     None.

Syntax:        parms:write

###

___pause_______________________________________________________________

Description: Use to display a message and wait for the user to press

a mouse key, the Escape key, or the Enter key.  This

command is useful when the user must complete some task

(for example, inserting a sample) before the macro can

continue.

Arguments:   (optional)  Text to display (string).  If no string is

specified, the message "Press ENTER or ESC to continue"

is displayed.

Example:     pause "Insert sample.  Press Enter key when ready."

/* Display message and wait for user response. */

scan  /* Begin sample data collection. */

###

___peakAt______________________________________________________________

Description: Use to get the x-axis and y-axis values for the peak

closest to the wavenumber you specify.

Arguments:   The wavenumber for which to get values (float).

Returns:     The x- and y-axis location (wavenumber and intensity)

of the closest peak (float).  The x-axis return value

is placed in the y register.  The y-axis return value

is placed in the z register.

Example:    load "pstyrene.abs"  /* Load file into memory. */

peakat 1600          /* Find peak nearest 1600. */

###

___peakPercent_________________________________________________________

Description: Use to report a specified y-axis value as a percentage

of the reference peak (established with the peakRef

command).  If verbose is set to "on", you can obtain a

peak percentage report for several wavenumbers.

Arguments:   Up to 19 wavenumbers for which you want to perform the

calculation (float).

Returns:     The y-axis percentage value for the specified wave-

number or for the last specified wavenumber if more

than one wavenumber is specified (float).  The return

value is also placed in the z register.

Example:     load "pstyrene.abs"  /* Load the file into memory. */

verbose "on"         /* Set verbose to "on." */

peakat 3000          /* Get nearest peak location. */

peakref y            /* Set the reference peak to be */

/* the peak at the wavenumber */

/* reported by the peakat */

/* command. */

peakpercent  2000 1943  /* Report the y-axis values */

/* that correspond to 2000 and */

/* 1943 wavenumbers as */

/* percentages of the peak */

/* specified with the peakref */

/* command. */

verbose "off"        /* Set verbose to "off." */

Also See:    peakRef

###

___peakRef_____________________________________________________________

Description: Use to specify a reference peak for the peak

percentage calculation.  (See the peakPercent command.)

Arguments:   The wavenumber location of the peak (float).

Returns:     The x- y-axis values (wavenumber and intensity) for the

peak (float).  The y-axis return value is placed in the

z register.  The x-axis return value is placed in the y

register.

Example:     load "pstyrene.abs"  /* Load the file into memory. */

peakat 3000          /* Retrieve the nearest peak. */

peakref y            /* Set the reference peak to be */

/* the peak at the wavenumber */

/* found by the peakat command. */

peakpercent 1943     /* Report the y-axis value for */

/* the peak at 1943 as a */

/* percentage of the peak */

/* specified with the peakref */

/* command. */

Also See:    peakPercent

###

___peaks_______________________________________________________________

Description: Use to generate a peak report, and annotate the peaks

for the current sample.

Arguments:   None.

Returns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:     load "pstyrene.abs"  /* Load file into memory. */

peaks                /* Generate a peak report */

/* and print it to the report */

/* display. */

###

___pickDir_____________________________________________________________

Description: Use to display a directory selection dialog.

Arguments:   (optional)  The starting directory you want the dialog

to show (string).  If no argument is given, then "."

(the current directory) is used.

Returns:     A directory name (drive:\directory) if the OK button is

selected.  NOTE: The trailing slash is not included in

the returned string.  An empty string is returned if

the Cancel button is selected.

Example:     pickdir "c:\winfirst\data"

###

___pickFile____________________________________________________________

Description: Use to display a list of all files in the current

directory that match the DOS filename with optional

wildcards (? and *).  With this list displayed, ApPro

waits for the user to select a listed file.

Arguments:   Any valid DOS filename with optional wildcards, which

include ? and * (string).

Returns:     The name of the file selected or an empty string if

nothing is selected.

Example:     s = pickfile "*.abs"  /* List all files with the */

/* .abs extension, and store */

/* the selected file in the s /

/* register. */

a = stringlength s    /* Place the length of the file */

/* name in s into a. */

if (a > 0)            /* If a file was selected... */

load s              /* Load the file.

###

___pickPoints__________________________________________________________

Description: Use to select points on the displayed data and place

those points into a user-defined array.  The points are

stored in the array as xy pairs (x value, then y

value).  For example, for the first selected point, the

x value is stored in array:data[1], the y value is

stored in array:data[2].  Similarly, for the nth point,

the x value is stored in array:data[ (n * 2) -1 ], and

the y value is stored in array:data[ n * 2 ].)  You

can configure this command several ways, as explained

in the arguments section.  After you define the array,

it is resized, set to 1 dimension, and set to a type of

float.

Arguments:   The user array to add points to (array), followed by

two optional arguments (integers).  (You can use the

first or both optional arguments.  But to use the

second argument, you must use the first argument.)  The

first optional argument specifies the number of points

you want to selected for the array.  The second

optional argument sets the type of data cursor to use

as you select points and also speifies how x and y

values are picked.

For the first argument, specify a positive number as

the number of points to pick or specify 0 to indicate

no limit.  When you specify a set number, the

pickPoints command returns after you select the

specified number of points.  When no limits is set,

choose OK on the displayed dialog to end the point

picking.

For the second argument (type of data cursor), allowed

values are 0, 1, 2, 3, 8, 9, 10, 11 or no argument.

Effects are as follows:

Use 0 to specify a free-moving data cursor.  The actual

cursor location is used to select the x and y values

entered into the array.

Use 1 to specify a free-moving data cursor and to force

the selected x value to match the actual x data value

in the sample array that is closest to the cursor

location.

Use 2 to specify a "sticky" data cursor and to force

the selected y value to match the actual y data value

in the sample array that is closest to the cursor

location.  (The cursor sticks to the data trace, to

force the cursor's y coordinate to match the sample

data y coordinate at the current wavenumber location.)

Use 3 to get the combined effects of using 1 and 2: a

sticky data cursor and force selection of actual x and

y values from sample array.

Use 8 to specify a free-moving cursor and no "rubber

band" effect (so that a temporary line is not drawn

between selected data points).

Use 9 to combine the effects of 1 and 8.

Use 10 to combine the effects of 8 and 2.

Use 11 to combine the effects of 8, 1 and 2 (8 and 3).

If no arguments are specified, default settings are

used.  The default number of points is no limit.  The

default data cursor is the sticky cursor.

Returns:     1 (int).  The return value is also placed in the z

register.

Example:     float array datapnts     /* Declare a float array to */

/* hold selected points. */

pickpoints datapnts 0 1  /* Pick any number of data */

/* points with a free- */

/* moving cursor. */

print "You picked these data points:" /* Show text */

/* in the report display. */

/* Execute loop statements starting with x equal to */

/* 1 and increment x by 2 (step 2) at each iteration */

/* of the loop. Stop executing the loop when the */

/* value of x is greater than or equal to the number */

/* of points in the array (datapnts:points). */

for x = 1 to dataPnts:points step 2  /* Loop. */

{           /* Mark the start of the loop. */

print datapnts:data[x]   " , "   datapnts:data[x + 1]

/* Print out the select points using */

/* this format: x value, y value. */

}           /* Mark the end of the loop. */

###

___pickString__________________________________________________________

Description: Use to display the contents of a text file one line

at a time similar to pickFile.  With this file is

displayed, ApPro waits for the user to select a line.

Arguments:   Any valid DOS filename, the title to display, 1 or 0

indicating sort or don't sort the list, 1 or 0

indicating autofit the column widths or not.

The title, sorting and autofit parameters are optional.

Default title="Pick an item"

Default sorting=1

Default autofit=0

Returns:     The text selected or an empty string if

nothing is selected.

Example:     s = pickString "c:\myfile.txt" "Pick a topic" 1 1

###

___pks_________________________________________________________________

Description: Pks is an object that contains properties and methods

you can set and execute to generate peak reports.

Properties:  Properties for the pks object are:

peakColumns    Use to specify the number of columns to use in the

peak report.

Arguments:     The word that identifies the number of colums to

use: One, Two, Three or Four (string).

Syntax:        pks:peakColumns = "three"

peakFilter     Use to specify the peak filtering method to use.

Arguments:     The name of the peak filter you want to use: None,

3point, 5point, or spline (string).

Syntax:        pks:peakFilter = "5point"

peakMethod     Use to specify the peak report method.

Arguments:     The name of the peak reporting method you want to

use: regression, derivative, lineshape, threshold,

or singles (string).

Syntax:        pks:peakMethod = "lineshape"

peakSensitivity Use to specify the number of standard deviations

a data point must be above the baseline to be

considered a peak.

Arguments:     Any value (int).

Syntax:        pks:peakSensitivity = 2

pkSort:        Use to specify the peak sorting method.

Arguments:     The name of the peak sorting method you want to

use: frequency or intensity (string).

Syntax:        pks:pkSort = "frequency"

Methods:     Methods for the pks object are:

pick:          Use to generate the peak report.

Arguments:     None.

Syntax:        pks:pick

readMethod:    Use to load a math method (configuration) and put

in place all math parameters it specifies,

including properties for the pks object.

Arguments:     (optional)  The name of the math method to load.

If no name is specified, the Load Math Setup

dialog appears.

Syntax:        pks:readmethod

saveMethod:    Use to save all current math parameters, including

pks object properties, as a math method

(configuration).

Arguments:     (optional)  The name to use for the math method

you save.  If no name is specified, the Save Math

Setup dailog appears.

Syntax:        pks:savemethod

###

___plot________________________________________________________________

Description: Use to plot the current data and/or text using the

specified plot method.

Arguments:   The name of the plot method to use.

Returns:     0 (int).

Example:     load "pstyrene.abs"  /* Load file into memory. */

plot "plot1.ini"     /* Plot using the plot method */

/* with plot1.ini as its name. */

###

___plsPredict__________________________________________________________

Description: Use to perform a partial least squares prediction.  The

prediction uses a calibration file, which you specify

as an argument to the plsPredict command.  (The

calibration file is created by the Mattson PLS

Quant software package.  Refer to your PLS Quant online

Help or the PLS Quant User's Reference Guide for

further information.)  The sample of unknown component

concentration, for which component concentration is

predicted, can be a data file you specify as an

argument, or the active sample data file.  Prediction

results are placed in a WinFIRST report window.

Arguments:   From one to 19 arguments: two file names and up to 15

registers or one file name and up to 18 registers.  The

first argument is required while all others are

optional.  The first argument is the name, including

directory path if needed, of the PLS Quant calibration

file to use to perform the prediction.  The second

argument is the name, including directory path, of the

data file for which component concentration is to be

predicted.  A register containing this information can

be used, but it must appear as the second argument.  If

you do not provide this second file name, either as a

file name or in a register, component concentration is

predicted for the active sample data file.  All

subsequent arguments are expected to be registers.

Predicted concentration results are placed in these

registers in the order (from right to left) they are

defined in the PLS method.  If the second argument is a

register and it does not contain a file name, it is

used to store results.  If there are extra registers

(i.e., more registers than components in the method),

the extra registers are ignored.  If there are fewer

registers than components in the method, as many

registers as possible are filled with results.

Returns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

###

___popRegisters_______________________________________________________

Description: Use to restore from the ApPro register stack all user

register values.  (User registers are a through z and A

through Z).  (Use the pushRegisters command to save

register values to the stack.)  The standard analogy

for a stack mechanism is a spring-mounted push-up stack

of plates in a buffet line.  The stack is a LIFO (last

in, first out) structure.  That is, the last plate, or

set of register values, put on the stack is the first

one taken off (when you use the popRegisters command).

The ApPro register stack used by the pushRegisters and

popRegisters commands has a depth of five.

The pushRegisters and popRegisters commands are

especially useful when one macro needs to call one or

more other macros, and the subsequent macros may alter

register values that the first macro will need when it

regains control.  In such a case, you can use

pushRegisters to save register values to the stack just

before calling a subsequent macro, and use popRegisters

to restore register values upon (or before) returning

control to the first macro.  The important thing to

remember is that every "push" requires a "pop" if the

values you expect are to be restored.  For instance:

Macro A:           stack: [empty]

pushregisters     stack: [A]

call macro B

Macro B:

pushregisters     stack: [BA]

call macro C

Macro C:

popregisters      stack: [A]

return to Macro A

Macro A:

popregisters      stack: [empty]

NOTE: For every pushregisters, there is a

popregisters.

Arguments:   None.

Returns:     1 for success, 0 for failure (int).

NOTE: The command may fail if there is not enough

memory or if the stack is full (for pushResgisters) or

empty (for popRegisters).

Example:     clrscr /* Clear the display screen. */

print "Testing push & pop registers."

/* Display this message. */

a = 1             /* Set register a to 1. */

b = 2.34          /* Set register b to 2.34. */

c = TESTING..."   /* Place string in register c. */

d = 2             /* Set register d to 2. */

print "a = " a "; b = " b "; c = " c "; d = " d

/* Print register values. */

d = pushregisters /* Push all registers and place */

/* results in register d. */

print "pushregisters = " d /* Print the contents */

/* of register d. */

print "a = " a "; b = " b "; c = " c "; d = " d

/* Print register values. */

zeroAllregisters  /* Set all register values to */

/* zero. */

print "a = " a "; b = " b "; c = " c "; d = " d

/* Print register values. */

d = 0             /* Set register d to zero. */

d = popregisters  /* Pop all registers and place */

/* result in register d. */

print "popregisters = " d /* Print the contents */

/* of register d. */

print "a = " a "; b = " b "; c = " c "; d = " d

/* Print register values. */

Also See:    popRegisters, zeroAllRegisters

###

___power_______________________________________________________________

Description: Use to perform exponential multiplication (i.e., raise

x to the power of y).  For example: 2 raised to the

power of 3 is equal to 2 x 2 x 2 = 8.

Arguments:   The base number upon which the exponent acts (float),

and the exponent (float).  Even though both values can

be floating point numbers, there are limitations:

- if the base number is negative, the exponent must be

an integer.

- if the base number is 0, the exponent must be greater

than 0.

Returns:     The calculated value (float).

Example:     print (power 2.5 3.5)  /* Raise 2.5 to the power of */

/* 3.5 and print result */

/* on report display. */

###

___precision___________________________________________________________

Description: Use to set the number of decimal places to report for

both x-axis and y-axis values.

Arguments:   The number of decimal places to report (int).

Returns:     The value provided as the number of values to report.

(A random value is placed in the z register.)

Example:     load "pstyrene.abs"  /* Load file into memory. */

precision 3          /* Set the number of decimal */

/* places to show when printing */

/* a report. */

peaks                /* Generate a peak report. */

Also See:    parms:xdigits, parms:xprecision, parms:ydigits,

parms:yprecision

###

___predictX____________________________________________________________

Description: Use to predict the x value given a specific y value.

(Use this command after establishing a data points

file and using the calibrate command to calibrate that

file.)  Also, if verbose is set to "on," the predictX

command creates a report of the predicted x value for

each of for several given y values.

Arguments:   Up to 19 y-axis values (float).

Returns:     The predicted value (float).  The return value is also

placed in the z register.  If more than one y value is

given, the predicted x value for the last y value is

returned.

Example:     calibrate  "pairs.dat"  /* Do a linear regression */

/* on the data points in */

/* the pairs.dat file. */

predictx 2.5            /* Based on the calibration */

/* results, predict x when */

/* 2.5 is the y value. */

Also See:    calibrate, predictY, quant

###

___predictY____________________________________________________________

Description: Use to predict the y value given a specific x value.

(Use this command after establishing a data points

file and using the calibrate command to calibrate that

file.)  Also, if verbose is set to "on," the predictY

command creates a report of the predicted y value

for each of several given x values.

Arguments:   Up to 19 x-axis values (float).

Returns:     The predicted value (float).  The return value is also

placed in the z register.  If more than one x value is

given, the predicted y value for the last x value is

returned.

Example:     calibrate  "pairs.dat"  /* Do a linear regression */

/* on the data points in */

/* the pairs.dat file. */

predicty 2.5            /* Based on the calibration */

/* results, predict y when */

/* 2.5 is the x value. */

Also See:    calibrate, predixtX, quant

###

___print_______________________________________________________________

Description: Use to display text in the WinFIRST report window.  The

arguments are concatenated into a text string and then

printed to the screen.  When a variable or register is

used as an argument, the value stored within that

variable or argument is printed.

Arguments:   Any literal value, register, or variable.  (Up to 19

arguments are allowed.)

Returns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:     calibrate  "pairs.dat"  /* Do a linear regression */

/* on the data points in */

/* the pairs.dat file. */

predictx 2.5            /* Based on the calibration */

/* results, predict x when */

/* 2.5 is the y value. */

print "When y = 2.5, the predicted x value is:  "  z

/* Print the results.  In this case, there are two */

/* arguments: a text string which is printed exactly */

/* as it appears when typed, followed by a register */

/* that, in this example, contains the results from */

/* the predictx command.  The register contents, not */

/* the letter z, are printed after the text string. */

###

___pushRegisters_______________________________________________________

Description: Use to save all values held in user registers (a

through z and A through Z) to a stack mechanism.  (Use

the popRegisters command to restore register values

placed on the stack.)  The standard analogy for a stack

mechanism is a spring-mounted push-up stack of plates

in a buffet line.  The stack is a LIFO (last in, first

out) structure.  That is, the last plate, or set of

register values, put on the stack is the first one

taken off (when you use the popRegisters command).  The

ApPro register stack used by the pushRegisters and

popRegisters commands has a depth of five.

The pushRegisters and popRegisters commands are

especially useful when one macro needs to call one or

more other macros, and the subsequent macros may alter

register values that the first macro will need when it

regains control.  In such a case, you can use

pushRegisters to save register values to the stack just

before calling a subsequent macro, and use popRegisters

to restore register values upon (or before) returning

control to the first macro.  The important thing to

remember is that every "push" requires a "pop" if the

values you expect are to be restored.  For instance:

Macro A:           stack: [empty]

pushregisters     stack: [A]

call macro B

Macro B:

pushregisters     stack: [BA]

call macro C

Macro C:

popregisters      stack: [A]

return to Macro A

Macro A:

popregisters      stack: [empty]

NOTE: For every pushregisters, there is a

popregisters.

Arguments:   None.

Returns:     1 for success, 0 for failure (int).

NOTE: The command may fail if there is not enough

memory or if the stack is full (for pushResgisters) or

empty (for popRegisters).

Example:     clrscr /* Clear the display screen. */

print "Testing push & pop registers."

/* Display this message. */

a = 1             /* Set register a to 1. */

b = 2.34          /* Set register b to 2.34. */

c = TESTING..."   /* Place string in register c. */

d = 2             /* Set register d to 2. */

print "a = " a "; b = " b "; c = " c "; d = " d

/* Print register values. */

d = pushregisters /* Push all registers and place */

/* results in register d. */

print "pushregisters = " d /* Print the contents */

/* of register d. */

print "a = " a "; b = " b "; c = " c "; d = " d

/* Print register values. */

zeroAllregisters  /* Set all register values to */

/* zero. */

print "a = " a "; b = " b "; c = " c "; d = " d

/* Print register values. */

d = 0             /* Set register d to zero. */

d = popregisters  /* Pop all registers and place */

/* result in register d. */

print "popregisters = " d /* Print the contents */

/* of register d. */

print "a = " a "; b = " b "; c = " c "; d = " d

/* Print register values. */

Also See:    popRegisters, zeroAllRegisters

###

___qqMeasure__________________________________________________________

Description: Similar to qqPredict, except that values reported are not

predicted concentrations, but are the measured Y values for the unknown.

The value returned depends upon the particular QuickQuant method used.

Also See:    qqPredict

###

___qqPredict__________________________________________________________

Description: Use to perform a prediction using Mattson QuickQuant.  The

prediction uses a calibration file, which you specify

as an argument to the qqPredict command.  (The

calibration file is created by the Mattson QuickQuant

software package.  Refer to your QuickQuant online

Help for further information.)  The sample of unknown

component concentration, for which component concentration is

predicted, can be a data file you specify as an

argument, or the active sample data file.  Prediction

results are placed in a WinFIRST report window.

Arguments:   From one to 19 arguments: two file names and up to 15

registers or one file name and up to 18 registers.  The

first argument is required while all others are

optional.  The first argument is the name, including

directory path if needed, of the QuickQuant calibration

file to use to perform the prediction.  The second

argument is the name, including directory path, of the

data file for which component concentration is to be

predicted.  A register containing this information can

be used, but it must appear as the second argument.  If

you do not provide this second file name, either as a

file name or in a register, component concentration is

predicted for the active sample data file.  All

subsequent arguments are expected to be registers.

Predicted concentration results are placed in these

registers in the order (from right to left) they are

defined in the QuickQuant method.  If the second argument is a

register and it does not contain a file name, it is

used to store results.  If there are extra registers

(i.e., more registers than components in the method),

the extra registers are ignored.  If there are fewer

registers than components in the method, as many

registers as possible are filled with results.

Returns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Also See:    qqMeasure

###

___quant_______________________________________________________________

Description: Quant is an object that contains properties and

methods you can set and execute to perform

quantitative analysis.

Properties:  Properties for the quant object are:

method:        Used to set the name of the current quant method

file that contains xy pairs (string).

NOTE: This file must be in ASCII format with one

xy pair per line.

NOTE: The method property for the quant object is

the same as the calibrate command.

Arguments:     The name of the file that contains the method

information (string).

Syntax:        quant:method = "filename"

Methods:     Methods for the quant object are:

correlation:   Use to return the (floating point) correlation

from the calibration.

Arguments:     None.

Syntax:        quant:correlation

intercept:     Use to return the (floating point) intercept from

the calibration.

Arguments:     None.

Syntax:        quant:intercept

predict:       Used to perform the quantitative analysis based

on the calibrated method.

NOTE: The predict method for the quant object is

the same as the predictY command.

The predict method returns the floating point

value predicted using the y = mx + b linear

relationship established by the calibration.

Arguments:     None.

Syntax:        quant:predict

Example:       /* Load a method and do a calibration: */

quant:method = "c:\first\data\quant\xy.dat"

/* Find an absorbance value of interest: */

p = peakAt 3100

/* Predict a concentration from the absorbance: */

c = quant:predict p

print "Concentration = " c

slope:         Use to return the (floating point) slope from the

calibration.

Arguments:     None.

Syntax:        quant:slope

stddev:        Use to return the (floating point) standard

deviation from the calibration.

Arguments:     None.

Syntax:        quant:stddev

Also See:    calibrate, predixtX, predictY

###

___quit________________________________________________________________

Description: Use to exit the WinFIRST software.

Arguments:   None.

Returns:     None.

Example:     quit  /* Exit the WinFIRST software. */

###

___ras2abs_____________________________________________________________

Description: Use to convert transmittance data (the current sample)

to absorbance data.

Arguments:   None.

Returns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:     load "trans.ras"  /* Load file into memory. */

ras2abs           /* Convert data to absorbance. */

show              /* Display the absorbance data. */

###

___ras2drt_____________________________________________________________

Description: Use to convert transmittance data (the current sample)

to diffuse reflectance data.

Arguments:   None.

Returns:     1 (int).  The return value

is also placed in the z register.

Example:     load "trans.ras"  /* Load file into memory. */

ras2drt           /* Convert data to diffuse */

/* reflectance. */

show              /* Display the results. */

###

___ratio_______________________________________________________________

Description: Use to ratio sample single beam data with another

single beam data file (typically a background file).

Arguments:   (optional)  The name of the background single beam

(string).  If a filename is not included, the current

background data is used.

Returns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:     load "pstyrene.sbm"  /* Load file into memory. */

ratio "sample.bkg"   /* Ratio the pstyrene.sbm file */

/* with the sample.bkg file. */

show                 /* Display the results. */

###

___readIniFile_________________________________________________________

Description: Use to read single lines from a WinFIRST "ini" file.

(All WinFIRST configuration files, including data

collection methods, display configurations, math

configurations, etc., are stored as ini files.)  All

ini files take the following general form:

in filename.ini:

[SECTION]

label = value

...

NOTE: You can disturb the operation of WinFIRST and of

Windows by using this command to make incorrect changes

to an ini file.  Use this command only with great care

and with a sound understanding of ini files and there

use.

Arguments:   Three: 1) the ini file filename, 2) the section heading

in the ini file, and 3) the label of the item within

the specified section for which you want to read the

value.

Returns:     For success, the value associated with the specified

label in the specified section.  For failure, the

string "FAIL" (when the specified label does not exist

in the specified section).

Example:     s = readinifile "c:\...\test.ini" "USER" "xPrecision"

if ((stringCompare s "FAIL") == 0)

{

print "READ xPrecision = " s

/* Some limit testing, to safeguard WinFIRST: */

a = string2int s

if (a < 7)

{

a = a + 1

}

else if (a > 2)

{

a = a - 1

}

else

{

a = 3

}

b = writeinifile "c:\..\test.ini" "USER" "xPrecision" a

if (b == 1)

{

s = readinifile "c:\...\test.ini" "USER" "xPrecision"

print "2nd READ xPRecision = " s

}

else

{

print "writeIniFile FAILED"

}

}

else

{

print "readIniFile FAILED"

}

/* Now test failure modes. */

print "readIniFile failure test:"

s = readinifile "c:\...\bogus\bogus.ini" "USER" "xPrecision"

print "readIniFile failure returns " s

b = writeinifile "c:|..\bogus\bogus.ini" "USER" "xPrecision" a

print "writeIniFile failure test"

print "writeIniFile failure returns " b

###

___reference___________________________________________________________

Description: Reference is an object that contains properties and

methods you can set and execute to perform various

operations with the R1 (first) reference.

Properties:  Properties for the reference object are:

data:          Use for direct access to the reference data.

Set Syntax:    To set an individual data point:

reference:data[point] = value

where, "point" is an index between 0 and

(reference:points - 1), inclusive

and "value" is the y value at the specified index

(float).

To set entire data array:

reference:data = userarray:data

where, "userarray" is a user-defined array.

Get Syntax:    To retrieve an individual data point:

r = reference:data[point]

where, "point" is an index between 0 and

(reference:points - 1), inclusive (float).

To retrieve entire data array:

userarray:data = reference:data

where, "userarray" is a user-defined array.

NOTE: userarray acquires every property of the

reference array.

firstx:        Use to get the starting wavenumber of the

reference spectrum.

Arguments:     None (retrievable only).

Get Syntax:    r = reference:lastx

Get Returns:   (float).

irType:        Use to get or set the data type of the reference

data.

Arguments:     An index to identify the data type, where:

0 = absorbance (.abs), 1 = transmittance (.ras),

2 = diffuse reflectance (.drt), 3 = sample single

beam (.sbm), 4 = background single beam .(bkg),

5 = sample interferogram (.igm), and

6 = background interferogram (.big).

Set Syntax:    reference:irtype = 0

Get Syntax:    r = reference:irtype

Get Returns:   (int).

lastx:         Use to get the ending wavenumber of the reference

spectrum.

Arguments:     None (retrievable only).

Get Syntax:    r = reference:lastx

Get Returns:   (float).

name:          Use to specify or get the reference file name.

Arguments:     If setting the name, specify the name (string).

Set Syntax:    reference:name = "name"

Get Syntax:    r = reference:name

Get Returns:   (string).

numrefs:       Use to return the number of reference spectra in

the data window.

Arguments:     None.

Syntax:        r = reference:numrefs

Returns:       (int).

points (rows): Use to get the number of data points.  (This is

the same as the number of rows in the data array).

Arguments:     None (retrievable only).

Get Syntax:    r = reference:points

Get Returns:   (int).

start:         Use to get the starting index of the data array.

Arguments:     None (retrievable only).

Get Syntax:    r = reference:start

Get Returns:   (int).

title:         Use to get or set the title for the current

reference data file.

Arguments:     If setting the title, specify the title (string).

Set Syntax:    reference:title = "title"

Get Syntax:    r = reference:title

Get Returns:   (string).

Methods:     Methods for the reference object are:

clear:         Use to clear the current reference.

Arguments:     None.

Syntax:        reference:clear

clearAll:      Use to clear all reference spectra from the

data window.

Arguments:     None.

Syntax:        reference:clearall

load:          Use to load the current reference file (name set

by reference:name).

Arguments:     None.

Syntax:        reference:load

overlay:       Use to overlay the references and the sample.

Arguments:     None.

Syntax:        reference:overlay

rolldown:      Use to roll the reference stack down one position.

(This method works the same as the related

WinFIRST toolbar icon.)

Arguments:     None.

Syntax:        reference:rolldown

rollup:        Use to roll the reference stack up one position.

(This method works the same as the related

WinFIRST toolbar icon.)

Arguments:     None.

Syntax:        reference:rollup

stack:         Use to stack the references and the sample.

Arguments:     None.

Syntax:        reference:stack

superimpose:   Use to superimpose the references and the sample.

Arguments:     None.

Syntax:        reference:superimpose

Also See:    loadRef

###

___reload______________________________________________________________

Description: Use to reload the current sample from disk.  All

unsaved data changes are discarded.

Arguments:   None.

Returns:     1 (int).  The return value is also placed in the z

register.

Example:     load "pstyrene.abs"  /* Load file into memory. */

abs2ras              /* Convert the data to */

/* transmittance. */

save "pstyren1.ras"  /* Save the transmittance data. */

reload               /* Load pstyrene.abs again. */

Also See:    load, sample, show

###

___rename______________________________________________________________

Description: Use to change the name of a file.  (After using this

command, the file no longer exists under its old name.)

Arguments:   The original name of the file (string), followed by the

name of the new file (string).

Returns:     1 for success, 0 for failure (int).  (A random value is

placed in the z register.)

Example:     rename "pstyrene.abs"  "pstyrene.bak"

/* Rename the file originally named pstyrene.abs */

/* as a file with pstyrene.bak as its name. */

Also See:    cd, copy, dir, erase, md

###

___report______________________________________________________________

Description: Report is an object that contains properties and

methods you can set and execute to perform various

operations with the active WinFIRST report window.

Properties:  Properties for the report object are:

iconize:       Use to minimize the active report window to an

icon.  (Same as the minimize method for the report

object.)

Arguments:     None.

Syntax:        report:iconize

maximize:      Use to maximize the active report window to full

size.

Arguments:     None.

Syntax:        report:maximize

minimize:      Use to minimize the active report window to an

icon.  (Same as the iconize method for the report

object.)

Arguments:     None.

Syntax:        report:minimize

name:          Use to set the name (filename) for the report.

Arguments:     The name of the file (string).

Syntax:        report:name = "filename"

nextWindow:    Use to make the next report window in the window

list the active window.

Arguments:     None.

Syntax:        report:nextwindow

numLines:      Use to return the integer number of lines in the

active report.

Arguments:     None (retrievable only).

Syntax:        report:numlines

numWindows:    Use to return the number of open report windows.

Arguments:     None.

Syntax:        report:numwindows

top:           Use to set and retrieve the line number of the

top of the report area.

Arguments:     (optional)  Line number to set to top of report.

Set Syntax:    report:top 12  /* Set line 12 to top of report. */

Get Syntax:    r = report:top  /* Put top line number in r. */

Get Returns:   (int).

Methods:     Methods for the report object are:

clear:         Use to clear the report area.

Arguments:     None.

Syntax:        report:clear

new:           Use to create a new report window.

Arguments:     None.

Syntax:        report:new

print:         Use to print the active report in its entirety.

WinFIRST prints using the default printer

specified for Windows.  If the printer driver is a

Windows 3.1 driver, the report is printed in

landscape mode.  If the printer driver is older,

the report is printed as specifed (by the Windows

Control Panel).

NOTE: This method for the report object provides a

"bare bones" means of printing reports.  A more

flexible method is to save the report to disk or

copy it to the Windows Clipboard as an ASCII text

file and then load or paste the report into your

favorite wordprocessor for formatting and layout

to suit your needs.

Arguments:     None.

Syntax:        report:print

read:          Use to read the report identified by the name

property (and stored on disk) into the active

WinFIRST report window.

Arguments:     None.

Syntax:        report:read

show:          Use to redraw the active report window.  If the

window is iconized, it is returned to its previous

size.  If the window is behind other windows, it

is brought to the top.

Arguments:     None.

Syntax:        report:show

toClipboard:   Use to place a copy of the report text in the

active WinFIRST report window on the Windows

Clipboard.

Arguments:     None.

Syntax:        report:toclipboard

write:         Use to write (save) the current report text in

the active WinFIRST report window area as a report

file.  The name set by the name property is

automatically assigned to report file that is

created.

Arguments:     None.

Syntax:        report:write

###

___rewind______________________________________________________________

Description: Use to move to the beginning of a text or binary file

that is open for read and write.

Arguments:   None.

Returns:     1 for success, 0 for failure (int).  (A random value is

placed in the z register.)

Example:     fopen 1 "binary.dat"  /* Open binary.dat as file 1. */

fwrite 1 2.5          /* Write 2.5 to open file. */

fwrite 1 2300         /* Write 2300 to open file. */

rewind 1              /* Move to top of file 1. */

fread 1 a             /* Read first value into the a */

/*  register. */

fread 1 b             /* Read next value into the b */

/* register. */

print  a "," b        /* Print the read values. */

fclose 1              /* Close file 1. */

Also See:    append, eof, exist, fclose, filepos, finput, fopen,

fprint, fread, freadByte, fseek, fwrite, fwriteByte

###

___rmsNoise____________________________________________________________

Description: Use to calculate the root mean squared noise of the

current spectrum within a specified wavenumber range.

If verbose is set to "on," this value is automatically

printed to the screen.

Arguments:   The starting and ending wavenumbers (float).

Returns:     The calculated value for success, 0 for failure

(float).  The return value is also placed in the z

register.

Example:     load "pstyrene.abs"  /* Load file into memory. */

rmsnoise  2000 2250  /* Find root mean squared noise */

/* between 2000 and 2250 cm-1. */

print z              /* Display the results. */

###

___run_________________________________________________________________

Description: Use to execute an ApPro macro.  You can use this

command from the command line or from within a macro.

NOTE: The run command clears from the symbol table only

what it adds to the symbol table.  Keep this in mind

when using the run command to use a user-defined

function.

NOTE: The difference between the run command and the

include command is that the run command lets you "pass"

command line arguments to the macro you execute.  That

is, you can place specific values on the same line with

the run command.  It is the programmer's responsibility

to retrieve these arguments from the command line

argument structure before using run to call another

file which also uses command line arguments.  Command

line arguments with the run command are optional.

Arguments:   The name of the macro file to execute (string),

followed by as many as 8 (optional) command line

arguments.  The command line arguments can be any

literal value, any register, and any variable.

Returns:     0 (int).  The return value is also placed in the z

register.

Example 1:   load "pstyrene.abs"  /* Load file into memory. */

run "myNeta.mac" 3000 2500 2000  /* Execute the */

/* macro, and pass it 3 command line arguments. */

Example 2:   /* "Analysis.mac" */

x = cmdargs:args[1]  /* Place the first command */

/* line argument in register x. */

y = cmdargs:args[2]  /* Place the second command */

/* line argument in register y. */

z = cmdargs:args[3]  /* Place the third command */

/* line argument in register z. */

neta x y z           /* Calculate net absorbance for */

/* the values stored in registers x, y, and z. */

###

___sadtler2matt________________________________________________________

Description: Use to a translate a copy of a file in the Sadtler

format (.isp or .irf) to a file in the Mattson

FIRST data file format.  The translated file becomes

the current sample.  The original Sadtler file that

is translated is not removed or otherwise altered by

the sadtler2matt command.  You can use the translated

data file in the same way you would any data file

created by WinFIRST.

Arguments:   The directory path and name of Sadtler file (string).

Returns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:     sadtler2matt "\sadtler\data\pstyrene"

/* Translate file and make it the */

/* sample data file. */

###

___sample______________________________________________________________

Description: Sample is an object that contains properties and

methods you can set and execute to perform various

operations on the current sample data.

Properties:  Properties for the sample object are:

audit:         Use to add text you specify (as an argument) to

the sample data file audit trail.  If no audit

trails exists, one is created.

Arguments:     The text string to add, placed within double

quote marks (string).

Syntax:        sample:audit = "Audit Trail Text"

data:          Use for direct access to the sample data array.

Set Syntax:    To set an individual data point:

sample:data[point] = value

where, "point" is an index between 0 and

(sample:points - 1), inclusive and "value" is the

y value at the specified index (float).

To set an entire data array:

sample:data = userarray:data

where, "userarray" is a user-defined array.

Get Syntax:    To retrieve an individual data point:

r = sample:data[point]

where, "point" is an index between 0 and

(sample:points - 1), inclusive.

To retrieve the entire sample data array:

userarray:data = sample:data

where, "userarray" is a user-defined array.

NOTE: Used this way, userarray acquires every

property of the sample data array.

firstx:        Use to get the starting wavenumber of the sample

spectrum.

Arguments:     None (retrievable only).

Get Syntax:    r = sample:lastx

Get Returns:   (float).

iconize:       Use to minimize the active data window to an icon.

(This property is the same as the minimize

property for the sample object.)

Arguments:     None.

Syntax:        sample:iconize

irType:        Use to get or set the data type of the sample

data.

Arguments:     An index to identify the data type, where:

0 = absorbance (.abs), 1 = transmittance (.ras),

2 = diffuse reflectance (.drt), 3 = sample single

beam (.sbm), 4 = background single beam .(bkg),

5 = sample interferogram (.igm), and

6 = background interferogram (.big).

Set Syntax:    sample:irtype = 0

Get Syntax:    r = sample:irtype

Get Returns:   (int).

lastx:         Use to get the ending wavenumber of the sample

spectrum.

Arguments:     None (retrievable only).

Get Syntax:    r = sample:lastx

Get Returns:   (float).

maximize:      Use to maximize the active data window to full

size.

Arguments:     None.

Syntax:        sample:maximize

minimize:      Use to minimize the active data window to an icon.

(This property is the same as the iconize property

for the sample object.)

Arguments:     None.

Syntax:        sample:minimize

Example:       if (sample:points)

{

a = sample:numwindows

for i = 1 to a step 1

{

sample:minimize

}

}

name:          Use to get or set the file name of the current

sample.

Arguments:     If setting the name, specify the name (string).

Set Syntax:    sample:name = "name"

Get Syntax:    r = sample:name

Get Returns:   (string).

nextWindow:    Use to make the next data window in the window

list the active window.

Arguments:     None.

Syntax:        sample:nextWindow

numWindows:    Use to return the number of data windows

that are currently open.  This method is useful

for looping through all current samples.

Arguments:     None.

Syntax:        sample:numWindows

Example:       if (sample:points)

{

a = sample:numwindows

sample:maximize

for i = 1 to a step 1

{

pause "Press Enter to view next sample"

sample:nextWindow

}

}

points (rows): Use to get the number of data points.  (This

value is the same as the number of rows in the

data array.)

Arguments:     None (retrievable only).

Get Syntax:    r = sample:points

Get Returns:   (int).

start:         Use to get the starting index of the data array.

Arguments:     None (retrievable only).

Get Syntax:    r = sample:start

Get Returns:   (int).

title:         Use to set or get the title of the current sample.

Arguments:     If setting the title, specify the title (string).

Set Syntax:    sample:title = "title"

Get Syntax:    r = sample:title

Get Returns:   (string).

timestamp:     Use to get sample data file acquisition timestamp.

Arguments:     None.

Get Syntax:    r = sample:timestamp

Get Returns:   (string).

xAxisLabel:    Use to set the x axis label for the current data

window.

NOTE: The display is not changed until you use the

show method for the sample object.

Arguments:     The label (string) to use.

Set Syntax:    sample:xaxislabel = "Concentration"

yAxisLabel:    Use to set the y axis label for the current data

window.

NOTE: The display is not changed until you use the

show method for the sample object.

Arguments:     The label (string) to use.

Set Syntax:    sample:yaxislabel = "Arbitrary Units"

Methods:     Methods for the sample object are:

clear:         Use to clear (unload) the current sample.

Arguments:     None.

Syntax:        sample:clear

copy:          Use to copy the current sample data to the

WinFIRST clipboard.

Arguments:     None.

Syntax:        sample:copy

exportBitmap:  Use to export the active sample data as a bitmap

graphic to the Windows Clipboard.

Arguments:     None.

Syntax:        sample:exportbitmap

exportMetafile: Use to export the active sample data as a

metafile graphic to the Windows Clipboard.

Arguments:     None.

Syntax:        sample:exportmetafile

info:          Use to create and display a sample data

information report for the active sample data

file.

Arguments:     None.

Syntax:        sample:info

load:          Use to load the specified file (name set by

sample:name).

Arguments:     None.

Syntax:        sample:load

pasteAsReference: Use to paste the data currently on the WinFIRST

clipboard to the active data dindow as a reference

data file.  If no data window is open, a new window

is opened and the clipboard data is pasted to that

window as sample data.

Arguments:     None.

Syntax:        sample:pasteAsReference

pasteAsSample: Use to paste the data currently on the WinFIRST

clipboard to the data window as the sample data

file.  If no data window is open, a new window

is opened and the clipboard data is pasted to

that window as sample data.

Arguments:     None.

Syntax:        sample:pasteAsSample

reload:        Use to reload the current sample.

Arguments:     None.

Syntax:        sample:reload

reportAudit:   Use to place the audit trail report for the

current sample data in a new report window.

(Same as reportAuditTrail.)

Arguments:     None.

Syntax:        sample:reportAudit

reportAuditTrail:  Use to place the audit trail report for the

current sample data in a new report window.

(Same as reportAudit.)

Arguments:     None.

Syntax:        sample:reportAudit

show:          Use to display, or redisplay (and redraw), the

current sample data file in the active WinFIRST

data window.  If the WinFIRST data window is

iconized, it is returned to its previous size.  If

the data window is behind other windows, it is

brought to the top.

Arguments:     None.

Syntax:        sample:show

toBackground:  Use to make the current sample the current

background data.

Arguments:     None.

Syntax:        sample:toBackground

toClipboard:   Use to place the current sample on the Windows

Clipboard.

Arguments:     None.

Syntax:        sample:toclipboard

Also See:    load, reload, show

###

___save________________________________________________________________

Description: Use to save the current sample data to disk.

Arguments:   (optional)  Name to assign to the saved file (string).

If you do not specify a name, the file is saved using

the current sample file name (and any existing file

with that same name is replaced by the saved data

file).

Returns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:     load "pstyrene.abs"  /* Load file into memory. */

abs2ras              /* Convert the data to */

/* transmittance. */

save "pstyren1.ras"  /* Save the transmittance data. */

###

___scale_______________________________________________________________

Description: Use to scale the sample data by a specified value.

Each data point is multiplied by the specified scaling

factor.

Arguments:   The number to use as the scaling factor (float).

Returns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:     load "pstyrene.abs"  /* Load the file into memory. */

scale 2.5            /* Scale data using 2.5 as */

/* the scaling factor. */

show                 /* Show the results. */

###

___scaleTo_____________________________________________________________

Description: Use to scale the sample data to a specific value.  The

maximum y value in the sample data is scaled to equal

the value you specify, and all other data values are

then scaled by that same amount.

Arguments:  The value to you want to scale to (float).

Returns:    1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:    load "pstyrene.abs"  /* Load file into memory. */

scaleto 3.4          /* Scale the maximum y value to */

/* make is equal 3.4 and scale */

/* all other values by that */

/* same amount. */

show                /* Display the results. */

###

___scaleToOne__________________________________________________________

Description: Use to scale the sample data between zero and one.

(The minimum data value is set to zero, the maximum

data value is set to one and all other values are

scaled in between.

Arguments:   None.

Returns:     1 (int).  The return value is also placed in the z

register.

Example:     load "tester.abs"  /* Load file into memory. */

scaletoOne         /* Scale the data between */

/* zero and one. */

show               /* Display the results. */

###

___scaleToPeak_________________________________________________________

Description: Use to scale all reference data files in the active

data window or data area to a sample data peak you

specify.

Arguments:   The wavenumber location of the sample data peak to

which you want to scale reference data files (int).

Returns:     1 (int).  The return value is also placed in the z

register.

Example:     load "tester.abs"  /* Load file into memory. */

scaletoPeak 3205.4  /* Scale all reference data */

/* files to the sample data peak */

/* at 3205 wavenumbers.*/

show                /* Display the results. */

###

___scan________________________________________________________________

Description: Use to collect sample data from the spectrometer.

Arguments:   None.

Returns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:     ask "Do you want to collect a sample?"

/* Ask question and wait for response. */

if (z == 1)  /* If answer is y, read next line. */

scan        /* Collect sample data. */

###

___search______________________________________________________________

Description: Use to perform a library search to find the best

matches for the sample spectrum.

Arguments:   (optional)  The name of each library you want to search

(string).

NOTE: Enclose each separate library name within a

separate set of double quotation marks.

If you do not specify a library name, the list of

last-used search libraries, as shown on the WinFIRST

Library Search dialog is used.

Returns:     1 (int).  The return value is also placed in the z

register.

Example:     load "tmp.abs"   /* Load file into memory. */

search "icon-1"  /* Compare the sample data with */

/* spectra in the icon-1 library */

/* to find the best matches. */

###

___setAdvancedMenus____________________________________________________

Description: Use to switch the WinFIRST menu system from simple

to advanced.

Arguments:   None.

Returns:     1 if simple menus are set, 0 if advanced menus are

set.

Example:     setAdvancedMenus

###

___setSimpleMenus______________________________________________________

Description: Use to switch the WinFIRST menu system from advanced

to simple.

Arguments:   None.

Returns:     1 if simple menus are set, 0 if advanced menus are

set.

Example:     setSimpleMenus

###

___setStringElement____________________________________________________

Description: Use to set an individual character within a string.

Arguments:   The string in which you want to set the character

(string), position of the character you want to set

(int), the value to replace the character at the

specified position (int or char).  If the value to put

in the string is expressed as an integer, that integer

also serves as the decimal representation for the ASCII

character.

Returns:    The new text string (string).

Example:    x = "This is a string"  /* Place text in x register. */

y = setStringElement x  14 'i'  /* Replace the 14th */

/* character in the string */

/* with the letter i. */

Also See:    buildString, execString, getStringElement, isNumber,

string2Float, string2Int, stringCompare, stringIndex,

stringLeft, stringLength, stringLower, stringMid,

stringRight, stringUpper

###

___shell_______________________________________________________________

Description: Use to create a DOS shell (command.com from Windows) in

which you can enter DOS commands.  This is a separate

instance of DOS, and all DOS commands are valid.  To

leave the DOS shell, type exit.

Arguments:   None.

Returns:     32 (int).  (A random value is placed in the z

register.)

Example:     shell  /* Start a DOS shell through Windows.  Enter */

/* DOS commands as needed.  Type exit to */

/* return to ApPro. */

Also See:    dos, execute, exit, launch, system

###

___show________________________________________________________________

Description: Use to display the data display area (and any loaded

data).

Arguments:   None.

Returns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:     load "pstyrene.abs"  /* Load file into memory. */

peaks                /* Generate and display a */

/* peak report. */

show                 /* Display the data again. */

Also See:    load, reload, sample

###

___showFile____________________________________________________________

Description: Use to display a text file in the report area.

Arguments:   The name of the file to display (string).

Returns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:     showfile "file1.dat"  /* Display the file.data text */

###

___sin_________________________________________________________________

Description: Use to compute the sine of the specified value.

Arguments:   Value (in radians) for which you want the sine (int or

float).

Returns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:     y = sin 0  /* Place the sine of 0 in y register. */

x = sin z  /* Place sine of value stored in z */

/* register into x register. */

Also See:    acos, asin, atan, cos, tan

###

___sleep_______________________________________________________________

Description: Use to stop a macro for as many seconds as you specify.

If verbose is set to �on,� a countdown is displayed in

the WinFIRST status line.  All WinFIRST and ApPro input

is also �put to sleep.�

Arguments:   The number of seconds to sleep (int).

Returns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:     sleep 5  /* Sleep for 5 seconds. */

###

___smooth______________________________________________________________

Description: Use to smooth data using the Savitsky-Golay method.

Arguments:   An odd number of smooth points from 3 to 25, inclusive

(int).

Returns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:     load "co.abs"  /* Load file into memory. */

smooth 9       /* Perform Savitsky-Golay smooth. */

show           /* Display the results. */

Also See:    boxcar, smth

###

___smth________________________________________________________________

Description: Smth is an object that contains properties and methods

you can set and execute to perform spectral data

smoothing.

Properties:  Properties for the smth object are:

smoothMethod:  Use to specify the smoothing method.

Arguments:     The name of the smoothing method to use: boxcar,

cubic, quadratic, quartic, or quintic (string).

Syntax:        smth:smoothmethod = "cubic"

smoothPoints:  Use to specify the number of smooth point to use.

Arguments:     An odd number within the legal range for the

specified smooth method: for boxcar, 3 to 25;

for cubic or quadratic, 5 to 25; and for quartic

or quintic, 7 to 25.  Specified even numbers are

forced to the next lower odd number.  Any number

outside the legal range is forced within range

(int).

Syntax:        smth:smoothpoints = 15

Methods:          Methods for the smth object are:

process:       Use to perform the smooth.

Arguments:     None.

Syntax:        smth:process

readMethod:    Use to load a math method (configuration), and

put in place the math parameters it specifies,

including smth object properties.

Arguments:     (optional)  The name of the math method to load.

If you specify no method, the Load Math Setup

dialog appears.

Syntax:        smth:readmethod

saveMethod:    Use to save a math method (configuration),

including all current smth object properties.

Arguments:     (optional)  The name of the math method to save.

If you specify no method, the Save Math Setup

dialog appears.

Syntax:        smth:savemethod

Also See:    boxcar, smooth

###

___spc2matt____________________________________________________________

Description: Use to translate a copy of a file in the Spectral

Calc (Grams) format to a file in the Mattson FIRST

data file format.  The translated file becomes the

current sample.  The original Spectral Calc file that

is translated is not removed or otherwise altered by

the spc2matt command.  You can use the translated data

file in the same way you would any data file created by

WinFIRST.

Arguments:   The directory path and name of Spectra Calc file

(string).

Returns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:     spc2matt "\scalc\data\pstyrene"

/* Translate file and make it the */

/* sample data file. */

###

___spc2matt____________________________________________________________

Description: Use to a translate a copy of a file in the Spectral

Calc (Grams) format to a file in the Mattson FIRST

data file format.  The translated file becomes the

current sample.  The original Spectral Calc file that

is translated is not removed or otherwise altered by

the spc2matt command.  You can use the translated data

file in the same way you would any data file created by

WinFIRST.

Arguments:   The directory path and name of Spectra Calc file

(string).

Returns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:     spc2matt "\scalc\data\pstyrene"

/* Translate file and make it the */

/* sample data file. */

Also See:    dx2matt, matt2dx, matt2nicDX, matt2nicSX, matt2omnic,

matt2pcir, matt2sadtler, matt2spc, matt2xy, nic2matt,

sadtler2matt, xy2matt

###

___sqrt________________________________________________________________

Description: Use to compute the square root of the specified value,

or the square root of the current sample if no value is

specified.

Arguments:   (optional)  The value for which to calculate square

root (int or float).  This specified value must be

positive.

Returns:     If you give no value, 1 is returned and the return

value is placed in the z register (int).  If you give

a value, the square root of the specified value is

returned (float).  If you give a negative value, a

nonsense value is returnd and 1 is placed in the z

register (int).  If you give 0 as a value, 0 is

returned (float) and the z register is not modified.

Example 1:   load "pstyrene.abs"  /* Load file into memory. */

sqrt                 /* Compute the square root */

/* of the sample data. */

show                 /* Display the results. */

Example 2:   sqrt 81  /* Compute the square root of 81. */

###

___square______________________________________________________________

Description: Use to compute the square of the specified value, or

the square of the current sample if no value is

specified.  When the data is squared, only positive

values are squared.

Arguments:   (optional)  The value to square (float).

Returns:     If you give no value, 1 is returned and the return

value is placed in the z register (int).  If you give a

negative value, a nonsense value is returnd and 1 is

placed in the z register (int).  If you give 0 as a

value, 0 is returned (float) and the z register is not

modified.

Example:     load "pstyrene.abs"  /* Load file into memory. */

square               /* Calculate the square of */

/* the sample data. */

show                 /* Display the results. */

Example 2:   square 2.5           /* Calculate square of 2.5. */

###

___stack_______________________________________________________________

Description: Use to display the current sample and reference data in

the stack display style.  (The current sample is

displayed at the bottom of the data display area.  The

R1, or first, reference appears just above sample data

with the R2 reference above R1, R3 above R2 and R4

above R3.)  If a file name is provided as an argument,

that data file is added to the reference stack as the

R1 reference.  If you add a new reference with this

command, the data is displayed.  Otherwise, the next

time you display data it is displayed using the stack

display style.

Arguments:   (optional)  The name of the file to add to the

reference stack (string).

Returns:     0 (int).  The return value is also placed in the z

register.

Example:     load "mix2.abs"  /* Load file into memory. */

stack "mix.abs"  /* Stack mix2.abs with mix.abs */

/* and display the results. */

Also See:    overlay, superimpose

###

___statusMsg___________________________________________________________

Description: Use to display a message at the ApPro command line.

After the message is displayed, the macro immediately

continues running.

Arguments:   Up to 19 valid ApPro values: register, variable, or

literal values.

Returns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:     fopen 1 "sample.txt"  /* Open sample.txt as file 1. */

fread 1 a             /* Read first value to the a */

/* register. */

eof 1                 /* Ask if at end of the file. */

if (z == 1)           /* If at end of file... */

statusMsg "At end of file... "  /* Show message. */

else                  /* If not at end of file... */

fread 1 b           /* Read second value. */

fclose 1              /* Close the file. */

###

___string2Float________________________________________________________

Description: Use to convert a text string to a floating point

number.

Arguments:   The text to convert (string).

Returns:     The floating point number that is expressed by the

specified text (float).

Example:     /* This reads text from a file, and converts the */

/* text into a floating point number: */

fopen 1 "testfile.dat"  /* Open testfile.dat file. */

finput 1 a              /* Read first line in file */

/* and store in register a. */

fclose 1                /* Close the file. */

f = string2float a      /* Convert text in the a */

/* register to a floating */

/* point number and store */

/* in the f register. */

print "The first value in testfile.dat is " f

/* Print the results in the */

/* WinFIRST report window. */

Also See:    buildString, execString, getStringElement, isNumber,

setStringElement, string2Int, stringCompare,

stringIndex, stringLeft, stringLength, stringLower,

stringMid, stringRight, stringUpper

###

___string2Int__________________________________________________________

Description: Use to convert a text string to an integer.

Arguments:   The text to convert (string).

Returns:     The integer that is expressed by the text (int).

Example:     /* This reads text from a file, and converts the */

/* text into an integer: */

fopen 1 "testfile.dat"  /* Open testfile.dat file. */

finput 1 a              /* Read first line in file */

/* and store in register a. */

fclose 1                /* Close the file. */

i = string2int a        /* Convert text in the a */

/* register to an integer */

/* and store the value in */

/* the i register. */

print "The first value in testfile.dat is " i

/* Print the results in the */

/* WinFIRST report window. */

Also See:    buildString, execString, getStringElement, isNumber,

setStringElement, string2Float, stringCompare,

stringIndex, stringLeft, stringLength, stringLower,

stringMid, stringRight, stringUpper

###

___stringCompare_______________________________________________________

Description: Use to tell if two text strings have the same content

(are equal to one another).

Arguments:   The two text values to compare (string).

Returns:     1 if the text strings are identical, 0 if not (int).

The return value is also placed in the z register.

Example:     enter n "Type name of file to load (or q to abort)"

/* Get the file name and place it */

/* in the n register. */

a = stringcompare n "q"  /* Compare text in register */

/* n with the string "q" and */

/* store the return value in */

/* the a register. */

if (a == 0)      /* If the strings are not equal ... */

load n           /* Load the file named in the n */

/* register. */

Also See:    buildString, execString, getStringElement, isNumber,

setStringElement, string2Float, string2Int,

stringIndex, stringLeft, stringLength, stringLower,

stringMid, stringRight, stringUpper

###

___stringIndex_________________________________________________________

Description: Use to find the string index (character position) for

a specified string within a string.

NOTE: This command is letter-case sensitive.

Arguments:   The string to search for (string), followed by the

string in which to search (string).

Returns:     The first index position of the located string or 0 if

the string was not located (int).  The return value is

also placed in the z register.

Example:     local int index  /* Declare local integer variable. */

index = stringIndex "Pro" "ApPro"

/* The value 3 is assigned to the index variable. */

Also See:    buildString, execString, getStringElement, isNumber,

setStringElement, string2Float, string2Int,

stringCompare, stringLeft, stringLength, stringLower,

stringMid, stringRight, stringUpper

###

___stringLeft__________________________________________________________

Description: Use to extract part of the left end of a text string.

The stringleft command builds a new text string by

extracting a specified number of characters, starting

from the leftmost character.

Arguments:   The string from which to extract text (string),

followed by the number of characters to extract (int).

Returns:     The extracted text (string).

Example:     /* This extracts the day from the result of the time */

/* command and sets the title of the loaded sample */

/* to the current day:*/

string day  /* Declare a string variable to hold the */

/* text that represents the current day. */

t = time    /* Get the date and time, and place them */

/* in the t register. */

day = stringleft t 10  /* Get the day (10 leftmost */

/* characters) from the string */

/* stored in the t register. /*

load "tmp.abs"  /* Load file into memory. */

title day       /* Set title for tmp.abs to match */

/* string stored in day variable. */

Also See:    buildString, execString, getStringElement, isNumber,

setStringElement, string2Float, string2Int,

stringCompare, stringIndex, stringLength, stringLower,

stringMid, stringRight, stringUpper

###

___stringLength________________________________________________________

Description: Use to get the length (in number of characters) of a

specified string.

Arguments:   The string whose length you want to get (string).

Returns:     The length of the string (int).

Example:     /* This asks the user to enter a file name to use */

/* to save newly acquired sample data: */

pause "Insert sample.  Press Enter key when ready."

/* Ask user to insert the sample */

/* and then continue when ready. */

scan           /* Collect a sample. */

enter n "Enter file name to use for this sample."

/* Ask user for new file name and */

/* store in register n. */

l = stringlength n  /* Get name length in register */

/* n and store in register l. */

/* (Eight characters is maximum */

/* name length DOS allows.) */

if ( l > 8 )        /* If length is more than 8... */

a = stringleft n 8  /* Set name to use equal first */

/* 8 characters only. */

save a              /* Save the new data using name */

/* stored in register a. */

Also See:    buildString, execString, getStringElement, isNumber,

setStringElement, string2Float, string2Int,

stringCompare, stringIndex, stringLeft, stringLower,

stringMid, stringRight, stringUpper

###

___stringLower_________________________________________________________

Description: Use to make all characters in a string lowercase.

Arguments:   The string upon which to act (string).

Returns:     The modified text (string).

Example:     s = stringLower "ApPro"

print s   /* The following string is printed: appro. */

Also See:    buildString, execString, getStringElement, isNumber,

setStringElement, string2Float, string2Int,

stringCompare, stringIndex, stringLeft, stringLength,

stringMid, stringRight, stringUpper

###

___stringMid___________________________________________________________

Description: Use to extract part of the middle of a text string.

The stringmid command builds a new text string by

extracting a specified number of characters, starting

at a specified character position.  The starting

position is specified by its distance (in characters)

from the left of the text string.

Arguments:   The string from which to extract the text (string),

followed by the starting position (int), followed by

the number of characters to extract (int).  This last

argument is optional; when you do not specify it, the

whole rest of the specified string is extracted.

Returns:     The extracted text (string).

Example:     /* This extracts the local time from the result of */

/* the time command and sets the title of the loaded */

/* sample to the current local time: */

string localtime  /* Declare a string variable to */

/* hold text that represents the */

/* current time. */

t = time          /* Get the date and time, and */

/* store in the t register. */

localtime = stringmid t 12 8  /* Get the 8 middle */

/* characters (time) */

/* from the string in */

/* the t register and */

/* store in localtime. */

load "tmp.abs"    /* Load file into memory. */

title localtime   /* Set title for tmp.abs to the */

/* string in localtime variable. */

Also See:    buildString, execString, getStringElement, isNumber,

setStringElement, string2Float, string2Int,

stringCompare, stringIndex, stringLeft, stringLength,

stringLower, stringRight, stringUpper

###

___stringRight_________________________________________________________

Description: Use to extract part of the right side of a text string.

The stringright command builds a new text string by

extracting the right part of a string starting at the

specified character position, and continuing to the end

of the string.  The specified starting position is

specified by its distance (in characters) from the

right of the text string.

Arguments:   The string from which to extract text (string),

followed by the starting character position (int).

Returns:     The extracted text (string).

Example:     /* This extracts the year from the result of the */

/* time command and sets the title of the loaded */

/* sample to the current year: */

string year  /* Declare a string variable to hold */

/* text that represents current year. */

t = time     /* Get the date and time, and store */

/* in the t register. */

year = stringright t 4  /* Get the 4 rightmost */

/* characters in the string held in the */

/* t register (year) and place in year. */

load "tmp.abs"  /* Load file into memory. */

title year   /* Set title for tmp.abs to string */

/* stored in year variable. */

Also See:    buildString, execString, getStringElement, isNumber,

setStringElement, string2Float, string2Int,

stringCompare, stringIndex, stringLeft, stringLength,

stringLower, stringMid, stringUpper

###

___stringUpper_________________________________________________________

Description: Use to make all characters in a string uppercase.

Arguments:   The string upon which to act (string).

Returns:     The modified text (string).

Example:     s = stringUpper "ApPro"

print s   /* The following string is printed: APPRO. */

Also See:    buildString, execString, getStringElement, isNumber,

setStringElement, string2Float, string2Int,

stringCompare, stringIndex, stringLeft, stringLength,

stringLower, stringMid, stringRight

###

___subtract____________________________________________________________

Description: Use to subtract reference data (of the same resolution)

from the current sample data.

Arguments:   (optional)  Name of the file to use as a reference

(string), followed by a subtraction factor (float).

You must supply a subtraction factor, but if you give

no file name, the subtract command uses the R1

reference.

Returns:     1 if a supplied reference name loads correctly, 0 if

not (int).  The return value is also placed in the z

register.

Example:     load "pstyrene.abs"  /* Load file into memory. */

subtract "h2O.abs" 0.5  /* Subtract 0.5 of h20.abs */

/* from sample data. */

show                    /* Display results. */

Also See:    add, average, divide, multiply

###

___superimpose_________________________________________________________

Description: Use to superimpose loaded references over sample data.

If a file name is specified, it is loaded as the R1

reference.

Arguments:   (optional)  The name of the new reference (string).

Returns:     0 (int).  The return value is also placed in the z

register.

Example:     load "mix.abs"          /* Load file into memory. */

superimpose "mix2.abs"  /* Load mix2.abs and */

/* superimpose on sample. */

Also See:    overlay, stack

###

___syntax_check________________________________________________________

Description: Use this special predefined variable to check the

syntax of a macro file instruction without actually

executing the instruction.

NOTE: The syntax_check variable is like a register in

that its assigned value is not cleared when the ApPro

symbol table is cleared.

Arguments:   One of the following two strings: on, to put the

syntax_check variable into use, or off, to take the

syntax_check variable out of use.

Returns:     None.

Example:     To turn use of the syntax_check variable on, type the

following instruction at the ApPro command line:

syntax_check = "on"

This instruction assigns a value to the syntax_check

variable that causes all instructions entered at the

ApPro command line, including the macro files you run,

to be evaluated for correct syntax, but not executed.

To turn use of the syntax_check variable off, type the

following instruction at the ApPro command line:

syntax_check = "off"

NOTE: You can also include the instruction that turns

syntax checking on and off in the macro files you

write.  But usually this is less convenient, as you

then need to edit the file to make it execute (or not

execute).

###

___system______________________________________________________________

Description: Use to run another DOS or Windows executable program

(application).  To return to WinFIRST, you must

terminate (quit) the other application.

Arguments:   The command line that starts the application (string).

Usually this consists of a program name and 0 or more

command line arguments.

Returns:     32 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:     system "pbrush.exe"  /* Run Windows Paintbrush, and */

/* wait for user to quit it. */

Also See:    dos, execute, exit, launch, shell

###

___tan_________________________________________________________________

Description: Use to compute the tangent of the specified value.

Arguments:   The value (in radians) for which to calculate the

tangent (int or float).

Returns:     Result (float).

Example:     y = tan 0  /* Place the tangent of 0 in register y. */

x = tan z  /* Place tangent of value in register z */

/* into register x. */

Also See:    acos, asin, atan, cos, sin

###

___tileWindows______________________________________________________

Description: Use to arrange all open data and report windows as

vertical or horizontal tiles.

Arguments:   None or 0 for vertical tiles. 1 for horizontal

tiles.

Returns:     1 for success, 0 for failure (int).

Example:     load pure.abs     /* Load data file. */

peaks             /* Generate peak report. */

load pure2.abs    /* Load data file. */

peaks             /* Generate peak report. */

load pstyrene.abs /* Load data file. */

peaks             /* Generate peak report. */

tilewindows 0     /* Arrange all open data and */

/* report windows as vertical */

/* tiles. */

Also See:    cascadeWindows, closeAllWindows

###

___time________________________________________________________________

Description: Use to get the system (computer) date and time.

Arguments:   None.

Returns:     The system date and time (string).  For example:

"Fri Jan 06 15:45:01 1995"

Example:     t = time    /* Get the system time and store in */

/* register t. */

print t     /* Print the system time. */

###

___timerCheck__________________________________________________________

Description: Use to find the number of seconds that have elapsed

since the timer was started, with the timerStart

command.

Arguments:   None.

Returns:     The number of seconds (int).  The return value is also

placed in the z register.

Example:     int secs  /* Declare an integer variable to hold */

/* the number of seconds. */

ask "Do you want to collect a sample?"

if (z == 1)  /* If answer is y... */

scan       /* Collect sample data. */

else         /* Otherwise... */

exit       /* Exit the macro. */

timerstart   /* Start the timer. */

enter n "Enter a name for this sample."

/* Ask for sample name and put in n. */

secs = timercheck  /* Get elapsed time and store */

/* in the secs variable. */

if (secs > 3600)   /* If time is more than 1 hour... */

{

enter a "This sample is too old.  Save it?"

*/ Ask this question. */

if (z == 1)        /* If answer is y... */

save  n          /* Save with name in n. */

}

else         /* Otherwise (time less than 1 hour)... */

save n     /* Save automatically with name in n. */

###

___timerStart__________________________________________________________

Description: Use to start the timer.  (Use the timerCheck command to

monitor the number of seconds elapsed since the timer

was started.)  Also, use the timerstart command to

reset (restart) the timer.

Arguments:   None.

Returns:     A relative system value.

Example:     int secs  /* Declare an integer variable to hold the */

/* number of elapsed seconds. */

ask "Do you want to collect a sample?"  /* Ask. */

if (z == 1)  /* If answer is y... */

scan        /* Collect sample data. */

else         /* Otherwise (answer is not y)... */

exit        /* Exit the macro. */

timerstart   /* Start the timer. */

enter n "Enter a name for this sample."

/* Ask for a file name for the sample. */

secs = timercheck  /* Get the elapsed time and store */

/* in the secs variable. */

if (secs > 3600)   /* If more than 1 hour elapsed... */

{

enter a "This sample is too old.  Save it?"

*/ Ask this question. */

if (z == 1)      /* If answer is y... */

save  n        /* Save sample with name in n. */

}

else         /* Otherwise (time less than 1 hour)... */

save n     /* Save automatically with name in n. */

###

___timeStampString_____________________________________________________

Description: Use to get the time stamp for the active sample data

file.

Arguments:   None.

Returns:     The sample data file's time stamp as text (string).

Example:     t = buildString timeStampString

/* Place the string returned by */

/* timeStampString into the t register. */

s = buildString "Unknown compound collected on" t

/* Place the new string returned by */

/* buildString into the s register. */

title s    /* Set the data file title to be the */

/* string in register s. */

###

___timeStampValue______________________________________________________

Description: Use to get the time stamp for the active sample data

and convert that value to a string.

Arguments:   None.

Returns:     The sample data file's time stamp (int).

Example:     int t1 load "pstyrene.abs"

t1 = timestampvalue

print t1

###

___title_______________________________________________________________

Description: Use to assign a title to the current sample.  (To save

the title with the sample data, use the save command.)

Arguments:   The title to be associated with the spectrum (string).

Returns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:     load "tmp.abs"     /* Load file into memory. */

title "sample1  Fri, Jan 1, 1993"  /* Assign a title */

/* to the sample. */

save "samp1a.abs"  /* Save sample (with title) as */

/* samp1a.abs. */

###

___truncate____________________________________________________________

Description: Use to extract a wavenumber (or data point) region from

the sample spectrum (or sample interferogram).

Arguments:   The starting and ending wavenumbers (or data points)

for the range you want to extract (float).

Returns:     1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:     load "pstyrene.abs"  /* Load file into memory. */

truncate 3000 2900   /* Cut off data outside 3000 */

/* and 2900 wavenumbers. */

show                 /* Display the resulting data. */

###

___type________________________________________________________________

Description: Use to return a string that indicates the type of an

ApPro expression.  NOTE: This command is useful for

debugging a macro when trying to determine the current

type of a register, since they are polymorphic (i.e.,

can be int, float, string), or the type of an

expression.  For example, type (1 + 1.0) returns

"FLOAT".

Arguments:   Any ApPro expression.

Returns:     Possible return values are: "STRING", "INTEGER",

"FLOAT", "ARRAY", "BYTE" and "CHAR" (string).

Example:     a = type (1 + 1.0);print a

###

___value_______________________________________________________________

Description: Use to get the y value for a specified wavenumber.

Also, if verbose is set to "on", the value command

supplies report of the y values for several given

wavenumbers.

Arguments:   Up to 19 wavenumbers for which you want y values

(float).

Returns:     The y value for the given wavenumber (float).  Also,

the return value is stored in z register.  (If more

than 1 wavenumber value is specified, the y value for

the last wavenumber is returned.)

Example:     load "pstyrene.abs"   /* Load file into memory. */

verbose "on"          /* Set verbose to "on." */

value 2000 3000 3500  /* Get report of y values. */

/*  Also, place y value for */

/* 3500 in the z register. */

verbose "off"         /* Set verbose to "off." */

###

___verbose_____________________________________________________________

Description:  Use to control the amount of information displayed by

certain commands, including p2pNoise, peakPercent,

predictX, predictY, rmsNoise, sleep, and value.

Arguments:    An "on" setting allows added information, any other

setting specifies minimal information (string).

Returns:      1 for success, 0 for failure (int).  The return value

is also placed in the z register.

Example:      load "pstyrene.abs"  /* Load the file into memory */

verbose "on"         /* Set verbose to "on." */

peakat 3000          /* Retrieve the nearest peak. */

peakref y            /* Set the reference peak to */

/* be the wavenumber retrieved */

/* by the peakat command. */

peakpercent  2000 1943  /* Report the y-axis values */

/* that correspond to 2000 and */

/* 1943 wavenumbers as */

/* percentages of the peak */

/* specified with the peakref */

/* command. */

verbose "off"       /* Set verbose to "off." */

###

___version_____________________________________________________________

Description: Use to get information for your current version of

WinFIRST software.

Arguments:   None.

Returns:     Version information (string).

Example:     z = version  /* Get version information and place */

/* it in the z register. */

print z      /* Print the information in z to the */

/* screen. */

###

___while_______________________________________________________________

Description: In general, use to test the truth of an expression,

and to then continue according to the results.  As

long as the expression is evaluated to true, the block

is executed.  After the "while" block is executed,

control returns to the "while" statement.  If the

expression evaluates to true again, the block is

executed again.  This continues until the expression

evaluates to false, or a loop break is used.

Example:     A = 0             /* Set A equal to 0/ */

while (A < 2)     /* While A is less than 2... */

begin             /* Start of "while" block. */

loadref         /* Load a reference. */

A = A + 1       /* Increment the value of A. */

end               /* End of "while" block. */

Also See:    break, choice, choice and branch, decision loop, for,

if

Refer also to the expanded discussion of while loops

and other control statements in Chapter 2 of "Using

WinFIRST Macros and ApPro."

###

___WriteIniFile________________________________________________________

Description: Use to write single lines to a WinFIRST "ini" file.

(All WinFIRST configuration files, including data

collection methods, display configurations, math

configurations, etc., are stored as ini files.)  All

ini files take the following general form:

in filename.ini:

[SECTION]

label = value

...

This command is particularly useful for changing ini

files on the fly and for changing settings for which

there is no macro or menu interface.

NOTE: You can disturb the operation of WinFIRST and of

Windows by using this command to make incorrect changes

to an ini file.  Use this command only with great care

and with a sound understanding of ini files and there

use.

Arguments:   Four: 1) the ini file filename, 2) the section heading

in the ini file, 3) the label of the item within the

specified section for which you want to write a value,

and 4) the value you want to write.

Returns:     1 for success, 0 for failure (int).

Example:     s = readinifile "c:\...\test.ini" "USER" "xPrecision"

if ((stringCompare s "FAIL") == 0)

{

print "READ xPrecision = " s

/* Some limit testing, to safeguard WinFIRST: */

a = string2int s

if (a < 7)

{

a = a + 1

}

else if (a > 2)

{

a = a - 1

}

else

{

a = 3

}

b = writeinifile "c:\..\test.ini" "USER" "xPrecision" a

if (b == 1)

{

s = readinifile "c:\...\test.ini" "USER" "xPrecision"

print "2nd READ xPRecision = " s

}

else

{

print "writeIniFile FAILED"

}

}

else

{

print "readIniFile FAILED"

}

/* Now test failure modes. */

print "readIniFile failure test:"

s = readinifile "c:\...\bogus\bogus.ini" "USER" "xPrecision"

print "readIniFile failure returns " s

b = writeinifile "c:|..\bogus\bogus.ini" "USER" "xPrecision" a

print "writeIniFile failure test"

print "writeIniFile failure returns " b

###

___xy2matt_____________________________________________________________

Description: Use to read a table of xy values, convert it to a

specified data type, and display it as the current

sample.  If a data type is not specified, ApPro prompts

for one.

Arguments:   The name of the file that contains the xy values

(string), followed by a number that represents a data

type (int), where 1 = abs, 2 = ras, 3 = drt, 4 = sbm,

5 = bkg, 6 = igm, and 7 = big.

Returns:     0 (int).  The return value is also placed in the z

register.

Example:     xy2matt pstyrene.tbl 1  /* Load table (pstyrene.tbl) */

/* as an absorbance file. */

Also See:    dx2matt, matt2dx, matt2nicDX, matt2nicSX, matt2omnic,

matt2pcir, matt2sadtler, matt2spc, matt2xy, nic2matt,

sadtler2matt, spc2matt

###

___zeroAllRegisters____________________________________________________

Description: Use to set the value of all regsiters to zero (0).  All

registers values are lost as they are replaced by zero.

Arguments:   None.

Returns:     0 (int).  The return value is also placed in the z

register.

Example:     zeroallregisters /* Set all register values to zero. */

Also See:    pushRegisters

###

___any expression_____________________________________________________

Description: Use to enter any ApPro built-in macro command.  You

should have a good working knowledge of ApPro

commands before using this advanced feature.

###