## Summary
Service Report
COMPANY |  |  | Cal Tech
CUSTOMER |  |  | Bruce Brunschwig
DATE |  |  | 2003-12-11
ENGINEER |  |  | Scott
Work completed
1. | Replaced Glassman HV
2. | replaced egun
3. | Brought up egun and aligned
4. | Ran Delta width and V1 calibartion
5. | Ran diag aquiitions to check counts and calibratiion
Follow-up Items
1. | replace loaner Glassman HV and egun
2.

## 8702
8702 PARAMETERS
COMPANY | Cal Tech
CUSTOMER | Bruce Brunschwig
DATE | 2003-12-11 |  | COMMENT
TEST ENGINEER | Scott
SPOT | CURRENT | PIERCE | F1 | SPOT L | CALC L | SPOT W | PHOTO e
OFF | 2.2
150 |  |  |  |  | 0
300 |  |  |  |  | 0
600 |  |  |  |  | 0
800 |  |  | 8580 |  | 0
FILAMENT/ CATHODE PARAMETERS
| I 2kV | E 2kV | K POWER |  | I FIL |  | E 10kV

## Xprobe
X-Probe Performance Specification
COMPANY | Cal Tech
CUSTOMER | Bruce Brunschwig
DATE | 2003-12-11 |  |  | COMMENT | Au 60%
TEST ENGINEER | Scott
Unscanned Au 4f 7/2 One minute acquistion
RESOLUTION | SPOT SIZE
| 150 micron |  | 300 micron |  | 600 micron |  | 800 micron
| SPEC | ACT | SPEC | ACT | SPEC | ACT | SPEC | ACT
RES 1 - 25eV
RESOLUTION | 0.69 | 0.73 | 0.75 |  | 0.95 |  | 1.1
AREA - K CNT | 10 | 9 | 55 |  | 100 |  | 140
RES 2 - 50eV
RESOLUTION | 0.85 |  | 0.9 | 9.85 | 1.1 |  | 1.2
AREA - K CNT | 30 |  | 150 | 64 | 260 |  | 335
RES 3 - 100eV
RESOLUTION | 1.25 |  | 1.3 |  | 1.4 | 1.24 | 1.5
AREA - K CNT | 50 |  | 250 |  | 410 | 160 | 510
RES 4 - 150eV
RESOLUTION | 1.6 |  | 1.65 |  | 1.7 |  | 1.8 | 1.6
AREA - K CNT | 60 |  | 280 |  | 470 |  | 660 | 220
Scanned Au 4f 7/2 one scan - Area in thousands counts
RESOLUTION | SCAN W | SPOT SIZE | FWHM | FWHM | AREA | AREA | PEAK BE | GAUSS %
| eV | MICRON | SPEC | ACTUAL | TYPICAL | ACTUAL | eV
RES 1 - 25eV
RES 2 - 50eV
RES 2 - 50eV
RES 4 - 150eV
RES 4 - 150eV

## Continuity
Analyzer Continuity

## Service Checklist
Maintenance Checks