Sample Loading/ Unloading

010/22/2018

Loading Samples

- Check no stub in XPS.

- Check Turbo on & back chamber at low pressure.

- open gate 1.

- turn off IG1.

- close gate 0.

- Turn Pfeiffer (Turbo) off.

- close gate 2.

- Wait for gate 2 to close.

- open gate 3.

- load stub using bottom notch.

- wait for Pfieffer to stop.

- hold door, close gate 3.

- open gate 2.

- Pfeiffer on.

- wait for back chamber  < 1 x 10-4 Torr.

- turn on IG1.

- open gate 0.

- close gate 1.

- wait >½ hour till IG1
< 5 x 10-8 Torr.

Moving sample to XPS Stage

- IG1 < 510-8..

- Raise stage (switch 7).

- Move stage to (0,0,0,0) to load stub.

- Lower stage (switch 7).

- Center stage is x and y

- Make sure transfer Arm is fully retracted.

- Open gate 5.

- Push sample with  transfer arm onto Stage.

- Unlock sample from arm.

- Fully retract transfer arm to LL.

- Close gate 5.

- Raise stage (switch 7).

Unloading Sample

- check IG1 < 5 x 10-8 Torr.

- Move stage to (0,0,0,0) to unload stub.

- lower stage (switch 7).

- Make sure transfer arm is Fully retracted.

- open gate 5.

- move stub from XPS stage to arm.

- move stub to load lock.

- Fully retract Transfer arm.

- close gate 5.

- open gate 1.

- close gate 0.

- turn off IG1.

- turn off Pfeiffer.

- close gate 2.

- Wait for gate 2 to close

- open gate 3.

- remove sample.

- wait for Pfeiffer to stop.

- hold door, close gate 3.

- open gate 2.

- start Pfeiffer.

- wait for P < 1x10-4 Torr.

- open gate 0.

- close gate 1,turn on IG1.