**Sample Loading/ Unloading**

010/22/2018

**Loading Samples**

1. Check no stub in XPS.
2. Check Turbo on & back chamber at low pressure.
3. open gate 1.
4. turn off IG1.
5. close gate 0.
6. Turn Pfeiffer (Turbo) off.
7. close gate 2.
8. Wait for gate 2 to close.
9. open gate 3.
10. load stub using bottom notch.
11. wait for Pfieffer to stop.
12. hold door, close gate 3.
13. open gate 2.
14. Pfeiffer on.
15. wait for back chamber < 1 x 10-4 Torr.
16. turn on IG1.
17. open gate 0.
18. close gate 1.
19. wait >½ hour till IG1
< 5 x 10-8 Torr.

**Moving sample to XPS Stage**

1. IG1 < 5×10-8..
2. Raise stage (switch 7).
3. Move stage to (0,0,0,0) to load stub.
4. Lower stage (switch 7).
5. Center stage is x and y
6. Make sure transfer Arm is fully retracted.
7. Open gate 5.
8. Push sample with transfer arm onto Stage.
9. Unlock sample from arm.
10. Fully retract transfer arm to LL.
11. Close gate 5.
12. Raise stage (switch 7).

**Unloading Sample**

1. check IG1 < 5 x 10-8 Torr.
2. Move stage to (0,0,0,0) to unload stub.
3. lower stage (switch 7).
4. Make sure transfer arm is Fully retracted.
5. open gate 5.
6. move stub from XPS stage to arm.
7. move stub to load lock.
8. Fully retract Transfer arm.
9. close gate 5.
10. open gate 1.
11. close gate 0.
12. turn off IG1.
13. turn off Pfeiffer.
14. close gate 2.
15. Wait for gate 2 to close
16. open gate 3.
17. remove sample.
18. wait for Pfeiffer to stop.
19. hold door, close gate 3.
20. open gate 2.
21. start Pfeiffer.
22. wait for P < 1x10-4 Torr.
23. open gate 0.
24. close gate 1,turn on IG1.