# Restart of MProbe For GLA’s only

# Mprobe Restart After Vacuum Loss

- Turn off all electronics, (2505 Memory Interface, Spectrometer Power Supply, Glassman High voltage, Service Physics X-Ray Gun Controller, nti sputter gun controller)

- Turn off alarm on interlock box near bottom of rack and set all three bypass switches on box up.

- Push reset button on interlock box

- Press reset on back of control box over XPS

- If the Neslab circulator is off Restart it.

- Check that the transfer arm is full withdrawn.

- If the Cryo’s are up to a temperature >80 K or the pressure is higher then 1 x 10-7 Torr.

  - Turn off cryo-pump compressor

  - If the Turbo pump is on turn it off and wait till it stops.

  - Turn on interlock override on control box.

  - Check gate valves to see that 0, 1, 2, 5, and 6 are open and 3 and 4 are closed

  - Turn on the turbo pump and wait till it is up to speed.

  - Wait till the vacuum is < 1  10-6 T, this will take 1 tor 2 hours or more.

- If the cryo compressor is off turn it on (you may need to turn the cryo switch off and then on).

- Make sure that the cryo pump is working

- After the cryo temperature is down below 25K (about 1 to 3 hours) close gate valves 1 and 5 (only 0, 2 and 5 are open).

- On the interlock box set the alarm and three bypass switches to down (on)

- Turn off interlock override.

- Continue to next section.

# Restarting M-Probe X-Ray Gun

- Check that the vacuum system is OK

- Gates 0, 2, and 6 should be open with all others closed

- Turn off interlock override on back of control box over XPS.

- Make sure the system power switch on the back is up

- Turn on X-ray gun boxes (

  - Turn on Glassman high voltage and press HV on,

  - Turn on 9603 X-ray gun spot size controller

  - Make sure the Interlock OK LED comes on otherwise you cannot get current from gun.

- If the vacuum has gotten above 1 x 10-7 Torr or system was vented:

  - Turn "ramp" knob on spot size controller to slowest, fully clock wise (service switch down)

  - Press "start filament" on 9600 followed by "HV on" on Glassman and wait for it to read 2kV on the spot size controller box - watch pressure (Filament on LED should turn on, Xfer OK LED on, Panel meter should read I FIL mode and go to ~1.2 A)

  - Set the "service" switch UP, keeping "ramp" knob on slowest setting (c.w.) and changing "stand by" to "operate" (takes 8 h, watch pressure)

- If the vacuum did not get above 1 x 10-7 Torr:

  - Turn "ramp" knob on spot size controller to fastest, counter clock wise (service switch down)

  - Press "start filament" followed by "HV on" and wait for it to read 2 KV on the spot size controller box - watch the pressure

  - Set the "service" switch up and turn "ramp" knob to fastest, then press "operate" (takes 2 h, watch pressure)

- Once 10kV reached, degas anode by starting X-ray gun on 100 micron spot and slowly increasing the spot size until largest spot does not raise pressure above 2 x 10-8 Torr

- Turn off X-rays

- Turn on spectrometer boxes (top 3, spectrometer power supply, flood gun, memory interface)

- Open software, check "X-ray gun operate" in ESCA control panel

- Switch spot size controller from "manual" to "computer"

- Degas flood gun - check “Flood Gun” box in ESCA control panel and increase energy to 5 eV, watch pressure and wait for at least 1 hour

- Turn down flood gun energy, uncheck flood gun box, and close ESCA control panel

- Instrument is ready to use

XPS maintenance schedule Kratos and MProbe

- Every week:

  - Refill water on Affinity and Neslab circulators

  - Check water level on Hawk chiller

  - Check N2 level on HREELS and order new cylinders

  - Replenish gloves, IPA, and KimWipe

- Every 6 months:

  - Change M-Probe pump oil (last 08/2019)

  - Change M-Probe anode (3 3/8" Cu gasket, Al/Cu anode, 6 Au 1/8" screws PRT-6520-003) (last 08/2019)

  - Change M-Probe Hawk circulator deionizer (Thermo Sci combined DEM/OXY cartridge D8809) and water filter (Hytrex cartridge filter PRT-6530-003) (last 08/2019)

  - Change the water and air filters for the MProbe located in the overhead rack.(changed 08/2019)

  - Change the oil in the MProbe Fore pump.

  - Change Kratos Affinity circulator filter (Pentek filtration polydepth filter cartridge PD-1-934) (last 2013-09-25)

- Every 36 months (last 2013-02-17):

  - Check and change Kratos deionizer cartridge (84-789), or when deionizer stays on

  - Change house water filter

- Every 60 months:

  - Change tip seals on Kratos (last 2013-04-19) and EELS (last 2010-07-01) scroll pumps

  - Change house air filter (Motor guard M-723 filter element) (last 2010-07-01)

# Contacts for M-Probe: Service Physics (Bend, Or) (541) 318-8688

Zack Mehl Zach@sphysics.com - New service Engineer, Mike’s replacement

Bob Chaney: bob@sphysics.com - the boss, good contact email or phone for software or in general x1

# Barbara Siordia: barb@sphysics.com - sends packages x2

Ruth Chaney: ruth@sphysics.com - Bob's wife, does POs 541 322 9405