Restart procedure for XPS after an interlock failure that shuts the whole instrument down.

- turn off alarm on interlock box near bottom of rack.

- push the reset button on interlock box.

- restart the circulating temperature control pump on Neslab in corner

- Turn off cryo-pump.

- press reset on back of box over xps

- check valves to see that the valves to allow the turbo to pump out load lock.  You may need to shut and open the valve once.

- Start the turbos

- Turn on the cryo-pump

- turn on the alarm of the interlock box.

- Wait a minute for the cryo to start and make sure that in the little window of the cryo under the load lock that there is movement.

- When the cryo shows that its temperature has gone down below about 25K you can open the valves for the cryo's to pump out their chambers.