XPS startup with 9600

- Turn on power to 9600

  - get interlock ok

  - and off on

- For new gun put switch in service position up

  - hit Start Filament: There are Three cycles that the 9600 should go through automatically:

    - Cycle 1: the filament current should go up to 4 mA and then down to low value ~0.6 and then slowly sink to ~1.4 mA over a 5 hours (max).  If you flip the switch to the normal position the cycles are much faster.  Note that V2KV and I2KV are off at this point.

    - Cycle 2: V2kV on.  I2KV will ramp up to 5 mA of a few hours;

    - Cycle 3:  Need to hit Operate and should start 3rd part of ramp, does not happen now some problem.  So for now need to turn on Glassman and M&W and hit operate.  M&W should come on and Flow ok.  Ramp up Glassman by hand slowly.  See old notes for timing.  Now ready to operate.

- Make sure that ‘turn X-ray gun on” is clicked in “ESCA control panel” on computer program.

- Now push computer control on 9600.