Service Physics Inc

20340 Empire Ave, Suite E4

Bend, OR 97701 USA

Phone 541-318-8688	Fax 541-322-0121

FILAMENT & CATHODE RAILURE REPORT

Measurements when Filament and/or Cathode are removed from service

Customer:  Caltech        Gun # I8856-LSG034

Date removed from Service 4/27/06

Date Installed

Filament S/N:    1071    shipper 11-29-05

Cathode S/N:

1. Following values taken with Glassman Voltage at zero (0)

I2KC (Voltage (TP3-TP3)/40 ) 6.05 mA drifted to 6.45  (6.5 mA (front panel meter);

IFil  (Voltage: (TP1-TP2)/.33) 1.58 A;  (1.6 A (fp metter)

E2KV:	2.190 Kv	  Line Frequence  50 Hz     60 Hz

Spot Size Settings	Pierce Voltage

off		0.926

150		0.397 kV

300		0.311 KV

600		0.240 kV

1000(800)		0.135 kV

2: Valuyes with Glassman at maximum setting (10 KV)

E10KV:  9.8 KV (fp metter)  9.9 using probe

Spot size Setting		F1	Glassman current

off		8.3 kV (7.60)		2.5 mA

150		8.8 kV (8.15)		4.0 mA

300		8.8 kV (8.18)		7.5 mA

600		8.8 kV (8.20)		12.5 mA

1000 (800)		8.8 kV (8.16)		22.6 mA

3: Resistance within one (1) minute of turning gun off

C-Fil

C-Pierce

Pierce-Fil

C-Ground