# Restart After Vacuum Loss

- Turn off electronics, (2505 Memory Interface, Spectrometer Power Supply, Glassman High voltage, Service Physics X-Ray Gun Controller, nti sputter gun controller)

- Turn off cryo-pump compressor

- Turn off alarm on interlock box near bottom of rack

- Restart the Neslab circulator

- Set all three bypass switches on interlock box (bottom of rack) up.

- Push reset button on interlock box

- Press reset on back of control box over XPS

- Turn on interlock override on control box.

- Check gate valves to see that 0, 1, 2, 5, and 6 are open and 3 and 4 are closed

- Turn on the turbo pump and wait till vacuum is < 1  10-4 T

- Turn on the cryo-pump compressor

- Mark sure that in the little window of the cryo pump is working

- After the cryo temperature is down below 25K (about 1 to 3 hours) close gate valves 1 and 5 (only 0, 2 and 5 are open).

- On the interlock box set the alarm and three bypass switches to down (on)

- Turn off interlock override.

- Continue to next section.

# Restarting M-Probe X-Ray Gun

- Check vacuum system OK

- Gates 0, 2, and 6 should be open with all others closed

- Turn off interlock override if it is on

- Make sure system power switch on the back is up

- Turn on X-ray gun boxes (bottom two – Turn on Glassman high voltage and press HV on, and turn on 9603 X-ray gun spot size controller)

- If the vacuum has gotten above 5 x 10-8 Torr or system was vented:

  - Turn "ramp" knob on spot size controller to slowest, fully clock wise (service switch in service position, up)

  - Press "start filament" followed by "HV on" and wait for it to read 2kV - watch pressure (Filament on LED should turn on, Xfer OK LED on, Panel meter should read I FIL mode and go to 0.65 A)

  - Set the "service" switch UP, keeping "ramp" knob on slowest setting (c.w.) and changing "stand by" to "operate" (takes 8 h, watch pressure)

- If the vacuum did not get above 5 x 10-8 Torr:

  - Turn "ramp" knob on spot size controller to fastest, counter clock wise (service switch up)

  - Press "start filament" followed by "HV on" and wait for it to read 2 KV - watch pressure

  - Set the "service" switch up and turn "ramp" knob to fastest, then press "operate" (takes 2 h, watch pressure)

- Once 10kV reached, degas anode by starting X-ray gun on 100 micron spot and slowly increasing the spot size until largest spot does not raise pressure above 2 x 10-8 Torr

- Turn off X-rays

- Turn on spectrometer boxes (top 3, spectrometer power supply, flood gun, memory interface)

- Open software, check "X-ray gun operate" in ESCA control panel

- Switch spot size controller from "manual" to "computer"

- Degas flood gun - check “Flood Gun” box in ESCA control panel and increase energy to 5 eV, watch pressure and wait for at least 1 hour

- Turn down flood gun energy, uncheck flood gun box, and close ESCA control panel

- Instrument is ready to use

XPS maintenance schedule Kratos and MProbe

Every week:

- Refill water on Affinity and Neslab circulators

- Check water level on Hawk chiller

- Check N2 level on HREELS and order new cylinders

- Replenish gloves, IPA, and KimWipe

Every 6 months:

- Change M-Probe pump oil (last 2013-06-10)

- Change M-Probe anode (3 3/8" Cu gasket, Al/Cu anode, 6 Au 1/8" screws PRT-6520-003) (last 2013-06-10)

- Change M-Probe Hawk chiller deionizer (Thermo Sci combined DEM/OXY cartridge D8809) and water filter (Hytrex cartridge filter PRT-6530-003) (last2013-06-10)

- Change Kratos Affinity circulator filter (Pentek filtration polydepth filter cartridge PD-1-934) (last 2013-09-25)

Every 36 months (last 2013-02-17):

- Check and change Kratos deionizer cartridge (84-789), or when deionizer stays on

- Change house water filter

Every 60 months:

- Change tip seals on Kratos (last 2013-04-19) and EELS (last 2010-07-01) scroll pumps

- Change house air filter (Motor guard M-723 filter element) (last 2010-07-01)

# Contacts for M-Probe: Service Physics (Bend, Or) (541) 318-8688

Mike Keenlyside: mike@sphysics.com - main contact, does site visits, very helpful by email or phone  3

Bob Chaney: bob@sphysics.com - the boss, good contact email or phone for software or in general x1

# Barbara Siordia: barb@sphysics.com - sends packages x2

Ruth Chaney: ruth@sphysics.com - Bob's wife, does POs 541 322 9405