Changing an Empty Dewar:

- On the Dewar that is empty:

  - Shut the valve the Dewar Gas valve.

  - Shut the 3 way valve on the supply lines.

  - Shut the needle valve after the pressure reducing valve.

- Remove the Pressure reducing valve from the Dewar

- Place the empty Dewar in the Hall

- Shut the needle valve of the other Dewar

- On the new full Dewar:

  - Attach the Pressure reducing valve to the new Dewar.

  - Open the Dewar Gas Valve

  - Open the Needle valve on the new Dewar.

  - Open the line valve to vent and flush the line for a 10 seconds.

  - Rotate the line valve 180 degrees to allow the new Dewar to supply gas to the system

  - Set the pressure reducing valve to 70 psi

- Open the Needle valve on the other Dewar.

- Check the both Dewar recirculatig valves are closed.

- If the second Dewar is less that half full call plant and order a new Dewar (the instructions are on the wall next to the phone).

N2 Supply system: